User Manual

SDS1000CML+/SDS1000DL+

Digital Oscilloscope

UM0101A-E01A

SIGLENT TECHNOLOGIES CO,.LTD

SIGLENT

Declaration

Copyright © by SIGLENT TECHNOLOGIES CO,.LTD. All rights reserved.

Contents in this Manual are not allowed to copy, extract and translate

before being allowed by SIGLENT.

SDS1000CML+/SDS1000DL+       User Manual  I

SIGLENT

Brief Introduction

Characteristic:

● The volume of the oscilloscope is cabinet and it is portable

● 7” Color TFT LCD display

● 2 channels, Bandwidth: 40MHz-150 MHz

● Single real-time sampling rate is: 1Gsa/s;

Equivalent sampling rate is 50GSa/s.

● Trigger types: Edge, Pulse, Video,Slope and Alternative

● Unique Digital Filter function and Waveform recorder function

● Auto measure thirty two parameters and support all measurement

function.

●Two groups’ reference waveforms and twenty groups’ capture

waveforms and twenty groups’ setups internal save/recall function

and USB flash drive save/recall function.

● Cursor types: Manual mode, Track mode and Auto mode.

● Channel waveform and its FFT waveform display on split screen.

● Waveform Intensity and Grid Brightness can be adjusted.

● Menu display in the form of pop-up that in order to convenience users to

use it.

● Rich Screen display styles: Classical, Modern, Tradition, Succinct.

● Multiple Language User Interface.

● Support Multilingual online help system

●Standard interface: USB Host, USB Device, LAN port

Standard Accessories:

● 1:1/10:1 probe (2 PCS)

● Power Cable that fits the standard of destination country

● Qualified Certification.

● CD (including EasyScopeX computer software system)

● Quick Start

● USB Cable

II SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

General Safety Summary

Carefully read the following safety precautions to avoid person injury and prevent

damage to the instrument and any products connected to it. To avoid potential

hazards, please use the instrument as specified.

Only qualified technician should perform service procedures

To Avoid Fire or Personal Injure

Use Proper Power Line

Use only the special power line of the instrument which approved by local state.

Ground the Instrument

The instrument grounds through the protective terra conductor of the power line.

To avoid electric shock, the ground conductor must be connected to the earth.

Make sure the instrument is grounded correctly before connect its input or output

terminals.

Connect the Signal Wire Correctly

The potential of the signal wire is equal to the earth, so do not connect the signal

wire to a high voltage. Do not touch the exposed contacts or components.

Look Over All Terminals’ Ratings

To avoid fire or electric shock, please look over all ratings and sign instruction of

the instrument. Before connecting the instrument, please read the manual

carefully to gain more information about the ratings.

Not Operate with Suspected Failures

If you suspect that there is a damage of the instrument, please let a qualified

service personnel check it.

Avoid Circuit or Wire Exposed Components Exposed

Do not touch exposed contacts or components when the power is on.

Do not operate in wet/damp conditions.

Do not operate in an explosive atmosphere.

Keep the surface of the instrument clean and dry.

SDS1000CML+/SDS1000DL+       User Manual    III

SIGLENT

If the equipment is used in a manner not specified by the manufacturer, the

protection provided by the equipment may be impaired.

This product has been tested to the requirements of CAN/CSA-C22.2 No.

61010-1, second edition, including Amendment 1, or a later version of the

same standard incorporating the same level of testing requirements.

Not to use the product for measurements within other measurement

categories, such as CAT II, CAT III, CAT IV.

Not to use the equipment for measurements on mains circuits, not to use

the equipment for measurements on voltage exceed the voltage range

describe in the manual.

Only probe assemblies which meet the manufacturer’s specifications shall

be used.

The Responsible body or operator should refer to the instruction manual to

preserve the protection afford by the equipment. If the equipment is used in

a manner not specified by the manufacturer, the protection provided by the

equipment may be impaired.

Any parts of the device and its accessories are not allowed to be changed or

replaced, other than authorized by the manufacturer of his agent.

IV SDS1000CML+/SDS1000DL+       User Manual

SIGLENT

Safety Terms and Symbols

Terms used on the instrument. Terms may appear on the instrument:

DANGER: Indicates an injury or hazard that may be immediately happen.

WARNING: Indicates an injury or hazard that may be not immediately

happen.

CAUTION: Indicates that a potential damage to the instrument or other

property might occur.

Symbols used on the instrument. Symbols may appear on the instrument:

Hazardous   Protective       Warning          Earth Ground     Power

Voltage    Earth Ground                                       Switch

SDS1000CML+/SDS1000DL+    User Manual   V

SIGLENT

Content

Brief Introduction ....................................................................................................II

Chapter 1 Accidence ..............................................................................................1

1.1 Accidence of Panel and Display Information ........................................................... 2

1.2 Function Checking ................................................................................................... 5

1.3      Probe .................................................................................................................. 7

1.3.1 Probe Safety ................................................................................................. 7

1.3.2 Probe Attenuation Setting .......................................................................... 7

1.3.3 Probe Compensation .................................................................................. 8

Chapter 2 Functions Instruction and Operation .....................................................9

2.1 Menu and Control Button ....................................................................................... 10

2.2 Connector ............................................................................................................... 12

2.3 Auto Setup .............................................................................................................. 13

2.4 Default Setup .......................................................................................................... 14

2.5 Universal Knob ...................................................................................................... 15

2.6 Vertical System....................................................................................................... 15

2.6.1 CH1, CH2 Channel .................................................................................... 16

2.6.2 Using Vertical “Position” Knob and “Volt/div” Knob .............................. 21

2.6.3 Math Functions........................................................................................... 21

2.6.4 Using Ref .................................................................................................... 27

2.7 Horizontal System .................................................................................................. 29

2.7.1 Horizontal Control Knob............................................................................ 30

2.7.2 Window Zone ............................................................................................. 30

2.8 Trigger System ....................................................................................................... 32

2.8.1 Signal Source ............................................................................................. 33

2.8.2 Trigger Type................................................................................................ 34

2.8.3 Coupling ...................................................................................................... 47

2.8.4 Position........................................................................................................ 47

2.8.5 Slope & Level ............................................................................................. 48

2.8.6 Trigger Holdoff........................................................................................ 49

2.9 Acquiring Signals System ...................................................................................... 50

2.10 Display System..................................................................................................... 56

2.10.1 X-Y Format ............................................................................................... 59

2.11 Measure System ................................................................................................... 60

2.11.1 Scale Measurement ................................................................................ 60

2.11.2 Cursor Measurement .............................................................................. 60

2.11.3 Auto Measurement .................................................................................. 66

2.12 Storage System ..................................................................................................... 72

2.13 Utility System....................................................................................................... 85

2.13.1 System Status .......................................................................................... 88

2.13.2 Language .................................................................................................. 88

2.13.3 Self Calibration......................................................................................... 89

2.13.4 Self Test ....................................................................................................... 90

2.13.5 Updating the System Software .............................................................. 92

2.13.6 Pass/Fail ................................................................................................... 92

2.13.7 Waveform Record.................................................................................... 96

2.13.8 Recorder ................................................................................................... 99

2.13.9 Remote Control...................................................................................... 102

2.14 Online Help Function ......................................................................................... 107

Chapter 3 Prompting Messages and Troubleshooting ...................................... 108

VI SDS1000CML+/SDS1000DL+                          User Manual

SIGLENT

3.1 Prompting Messages: ........................................................................................... 108

3.2 Troubleshooting.................................................................................................... 110

Chapter 4 Service and Support .......................................................................... 112

4.1 Maintain Summary ............................................................................................... 112

4.2 Contact SIGLENT ................................................................................................ 113

Appendix A: Default Setup ................................................................................. 114

Appendix B: Daily Maintain and Cleaning .......................................................... 116

SDS1000CML+/SDS1000DL+                                   User Manual           VII

SIGLENT

Chapter 1 Accidence

SDS1000CML+ /SDS1000DL+ Series Digital Oscilloscope is mini-type and

portable bench type instruments, which could be used for measuring as the GND

voltage.

This Chapter shows you how to operate following tasks:

◆ Accidence of panel and Display information

◆ Simple checking of functions

◆ Matching probes attenuation coefficient

◆ Probe compensation

SDS1000CML+/SDS1000DL+     User Manual   1

SIGLENT

1.1 Accidence of Panel and Display Information

1.1.1 Front Panel

It is important for you to understand the DSO’s front panel before operating it. The

following contents are the brief introduction for the front panel function, which is

useful to be familiar with the operation of the SDS1000CML+/SDS1000DL+

Series Digital Storage Oscilloscope in short time.

The oscilloscopes provides an easy-to-use front panel to convenience users to

operate them, the panel contains knobs and buttons. There is a list of five ashen

buttons as menu operational buttons on the right of display screen. You can set

different options of the current menu in virtue of them. Other buttons are function

buttons; you can enter different function menus or obtain given function

application in virtue of them.

Picture 1.1-1 SDS1000CML+ Series Front Panel

2   SDS1000CML+/SDS1000DL+         User Manual

SIGLENT

No.     Description                  No.    Description

1     Power button                   10   Trigger Control Area

2     Menu On/Off                    11   Probe Compensation

3     Universal Knob                 12   Horizontal Control Area

4     Functions Menus                13   Ext Trigger Terminal

5     Default Setup                  14   Vertical Control Area

6     Help button                    15   Channel Input Terminal

7     Single Trigger                 16   Print key

8     Run/Stop Control               17   Menu Softkey

9     Auto Setup                     18   USB Host

1.1.2 Back and Side Connections

The following images show back and side panel connection locations.

Picture 1.1-2 SDS1000CML+ Back and Side panel

1. Handle

2. AC Power Input Terminal

3. USB Device Connector

4. Pass/Fail Output Connector

5. LAN Port

6. Lock Hole

SDS1000CML+/SDS1000DL+     User Manual 3

SIGLENT

1.1.3 User display interface

Picture 1.1-3

1.Product Logo

Siglent is the registered trademark of our company.

2. Trigger status

Armed. The oscilloscope is acquiring pre-trigger data. All triggers are ignored

in this state.

Ready. All pre-trigger data has been acquired and the oscilloscope is ready to

accept a trigger.

Trig’d. The oscilloscope has seen a trigger and is acquiring the posttrigger data.

Stop. The oscilloscope has stopped acquiring waveform data.

Auto. The oscilloscope is in auto mode and is acquiring waveforms in the

absence of triggers.

Scan. The oscilloscope is acquiring and displaying waveform data continuously in

scan mode.

3. USB Host connected mark.

4. Waveform memory

Show the position of the current waveform in the memory of the

oscillsocpe.

5. Trigger position.

Turn the HORIZONTAL POSITION knob to adjust the trigger position of the

4   SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

waveform.

6. Show the LAN port.

Indicates the LAN port is connected.

Indicates the LAN port is disconnected.

7. Show the Channel symbol.

8. Readout shows trigger signal frequency..

9. Readout shows the trigger level value and trigger type..

10.Readout shows the trigger delay of waveform.

11. Readout shows the main time base setting.

12. Icon shows the channel setting.

13. Icon shows the channel offset position.

14. Icon shows the trigger level position

1.2 Function Checking

When you check whether or not the oscilloscope could work smoothly, please

operate as following:

1. Power On the oscilloscope.

Press “DEFAULT SETUP” to show the result of the self check. The probe default

attenuation is 1X.

Picture 1.2- 1

2. Set the switch to 1X on the probe and connect the probe to channel 1 on the

oscilloscope. To do this, align the slot in the probe connector with the key on the

CH 1 BNC, push to connect, and twist to the right to lock the probe in place.

Connect the probe tip and reference lead to the PROBE COMP connectors

SDS1000CML+/SDS1000DL+      User Manual  5

SIGLENT

Picture 1.2-2

3.Press “AUTO” to show the 1 KHz frequency and about 3V peak-peak square

wave in couple seconds

Picture 1.2-3

4. Press “CH1” two times to cancel the channel 1, Press“CH2” to change screen

into channel 2, reset the channel 2 as step 2 and step 3.

6   SDS1000CML+/SDS1000DL+      User Manual

SIGLENT

1.3 Probe

1.3.1 Probe Safety

A guard around the probe body provides a finger barrier for protection from

electric shock.

Picture 1.3-1

Connect the probe to the oscilloscope and connect the ground terminal to ground

before you take any measurements.

Note: ● To avoid electric shock when using the probe, keep fingers behind the

guard on the probe body.

● To avoid electric shock while using the probe, do not touch metallic

portions of the probe head while it is connected to a voltage source.

Connect the probe to the oscilloscope and connect the ground terminal

to ground before you take any measurements.

1.3.2 Probe Attenuation Setting

Probes are available with various attenuation factors which affect the vertical

scale of the signal. The Probe Check function verifies that the Probe attenuation

option matches the attenuation of the probe.

You can push a vertical menu button (such as the CH 1 MENU button), and select

the Probe option that matches the attenuation factor of your probe.

Note: The default setting for the Probe option is 1X.

Be sure that the attenuation switch on the probe matches the Probe option in the

oscilloscope. Switch settings are 1X and 10X.

Note: When the attenuation switch is set to 1X, the probe limits the bandwidth of the

oscilloscope to 6MHz (according to Probe spec). To use the full bandwidth of the

oscilloscope, be sure to set the switch to 10X

SDS1000CML+/SDS1000DL+        User Manual     7

SIGLENT

1.3.3 Probe Compensation

As an alternative method to Probe Check, you can manually perform this

adjustment to match your probe to the input channel.

Picture1.3-2

1. Set the Probe option attenuation in the channel menu to 10X. Set the switch to

10X on the probe and connect the probe to channel 1 on the oscilloscope. If you

use the probe hook-tip, ensure a proper connection by firmly inserting the tip

onto the probe.

2. Attach the probe tip to the PROBE COMP~3V connector and the reference

lead to the PROBE COMP Ground connector. Display the channel and then

push the “AUTO” button.

3. Check the shape of the displayed waveform.

Over              Compensated            under

Compensated             correctly       compensated

Picture 1.3-3

4. If necessary, adjust your probe. Repeat as necessary.

8   SDS1000CML+/SDS1000DL+       User Manual

SIGLENT

Chapter 2 Functions Instruction and Operation

To use your oscilloscope effectively, you need to learn about the following

oscilloscope functions:

◆ Menu and control button

◆ Connector

◆ Auto Setup

◆ Default Setup

◆ Universal knob

◆ Vertical System

◆ Horizontal System

◆ Trigger System

◆ Acquiring signals System

◆ Display System

◆ Measuring waveforms System

◆ Utility System

◆ Storage System

◆ Online Help function

SDS1000CML+/SDS1000DL+       User Manual   9

SIGLENT

2.1 Menu and Control Button

Showing as the following picture:

Picture 2-1

■ Channel buttons (1, 2): Press a channel button to turn that channel ON or OFF

and open the channel menu for that channel. You can use the channel menu to

set up a channel. When the channel is on, the channel button is lit.

■ MATH: Press to display the Math menu. You can use the MAH menu to use the

oscilloscopes Math functions.

■ REF: Press to display the Ref Wave menu. You can use this menu to save and

recall four or two reference waveforms internal memory.

■ HORI MENU: Press to display the Horizontal menu. You can use the Horizontal

menu to display the waveform and zoom in a segment of a waveform.

■ TRIG MENU: Press to display the Trigger menu. You can use the Trigger menu

to set the trigger type (Edge. Pulse, Video, Slope, Alternative) and trigger

settings.

■ SET TO 50%: Press to stabilize a waveform quickly. The oscilloscope can

set the trigger level to be halfway between the minimum and maximum

voltage level automatically. This is useful when you connect a signal to the

EXT TRIG connector and set the trigger source to Ext or Ext/5.

10   SDS1000CML+/SDS1000DL+         User Manual

SIGLENT

■ FORCE: Use the FORCE button to complete the current waveform acquisition

whether the oscilloscope detects a trigger or not. This is useful for Single

acquisitions and Normal trigger mode.

■ SAVE/RECALL: Press to display the Save/Recall menu. You can use the

Save/Recall menu to save and recall up to 20 oscilloscope setups or waveforms

in internal memory (up to 20 waveforms) or on a USB memory device (limited by

memory capacity of USB device). You can also use it to recall the default factory

settings, to save waveform data as a comma-delimited file (.CSV), and to save

or print the displayed waveform image.

■ ACQUIRE: Press to display Acquire menu. You can use the Acquire menu to

set the acquisition Sampling Mode (Sampling, Peak Detect, Average).

■ MEASURE: Press to display a menu of measurement parameters.

■ CURSORS: Display the Cursor Menu. Vertical Position controls adjust cursor

position while displaying the Cursor Menu and the cursors are activated.

Cursors remain displayed (unless the “Type” option is set to “Off”) after leaving

the Cursor Menu but are not adjustable.

■ DISPLAY: Press to open the Display menu. You can use the Display menu to

set grid and waveform display styles, and persistence.

■ UTILITY: Press to open the Utlity menu. You can use the Utility menu to

configure oscilloscope features, such as sound, language, counter, etc. You can

also view system status and update software.

■ DEFAULT SETUP: Press to reset the oscilloscope’s settings to the default

factory configuration.

■ HELP: Enter the online help system.

■ AUTO: Automatically sets the oscilloscope controls to produce a usable display

of the input signals.

■ RUN/STOP: Continuously acquires waveforms or stops the acquisition.

Note:If waveform acquisition is stopped (using the RUN/STOP or SINGLE

button), the SEC/DIV control expands or compresses the waveform.

■ SINGLE: Acquire a single waveform and then stops.

SDS1000CML+/SDS1000DL+         User Manual   11

SIGLENT

2.2 Connector

Picture 2-2

■ Channel Connector (CH1, CH2): Input connectors for waveforms display.

■ EXT TRIG: Input connector for an external trigger source. Use the Trigger

Menu to select the “Ext” or “Ext/5” trigger source.

■ Probe Component: Voltage probe compensation output and ground. Use

to electrically match the probe to the oscilloscope input circuit.

Note:If you connect a voltage source to a ground terminal, you

may damage the oscilloscope or the circuit under test. To avoid

this, do not connect a voltage source to any ground terminals.

12   SDS1000CML+/SDS1000DL+            User Manual

SIGLENT

2.3 Auto Setup

The SDS1000CML+/SDS1000DL+ Series Digital Storage Oscilloscopes have a

Auto Setup function that identifies the waveform types and automatically adjusts

controls to produce a usable display of the input signal.

Press the AUTO button, and then press the menu option button adjacent to the

desired waveform as follows:

Picture 2-3

Table 2-1 Auto Set function Menu:

Option                         Description

(Multi-cycle sine) Auto set the screen and display several

cyc signal.

(Single-cycle sine) Set the screen and auto display single cyc

signal.

(Rising edge)           Auto set and show the rising time.

(Falling edge)            Auto set and show the falling time.

(Undo Setup)            Causes the oscilloscope to recall the

previous setup.

Auto set determines the trigger source based on the following conditions:

● If multiple channels have signals, channel with the lowest frequency signal.

● No signals found, the lowest-numbered channel displayed when Auto set was

invoked

● No signals found and no channels displayed, oscilloscope displays and uses

channel 1.

SDS1000CML+/SDS1000DL+       User Manual 13

SIGLENT

Table 2-2     Auto set the function item

Function                    Setting

Acquire Mode                Adjusted to Sampling

Display Format              Y-T

Display Type                Set to Dots for a video signal, set to Vectors

for an FFT spectrum; otherwise, unchanged

Vertical Coupling           Adjusted to DC or AC according to the input

signal

Bandwidth Limit             Off(full)

V/div                       Adjusted

VOLTS/DIV adjustability     Coarse

Signal inverted             Off

Horizontal position         Center

S/div                       Adjusted

Trigger type                Edge

Trigger source              Auto detect the channel which has the input

signal

Trigger slope               Rising

Trigger mode                Auto

Trigger coupling            DC

Trigger holdoff             Minimum

Trigger level               Set to 50%

2.4 Default Setup

The oscilloscope is set up for normal operation when it is shipped from the factory.

This is the default setup. To recall this setup, press the DEFAULT SETUP button.

The options, buttons and controls that change settings when you press the

DEFAULT SETUP button, refer to appendix B.

The DEFAULT SETUP button does not reset the following settings:

● Language option

● Saved reference waveform files

● Saved setup files

● Display contrast

● Calibration data

14  SDS1000CML+/SDS1000DL+         User Manual

SIGLENT

2.5 Universal Knob

Picture 2-5 Universal Knob

You can use the Universal knob with many functions, such as adjusting the

holdoff time, moving cursors, setting the pulse width, Setting the Video Linage,

adjusting the upper and lower frequency limit, adjust X and Y masks when

using the pass/fail function etc. You can also turn the “Universal” knob to

adjust    the  storage   position   of  setups,    waveforms,    pictures    when

saving/recalling and to select menu options.

2.6 Vertical System

The vertical control could be used for displaying waveform, rectify scale and

position.

Volt/div Knob

Vertical Position

Knob

Picture 2.6-1

SDS1000CML+/SDS1000DL+        User Manual       15

SIGLENT

2.6.1 CH1, CH2 Channel

Table 2-1 CH1, CH2 function menu 1:

Option         Setting        Introduction

DC             DC passes both AC and DC components of

the input signal.

Coupling       AC             AC blocks the DC component of the input

signal and attenuates signals below 10 Hz.

GND            GND disconnects the input signal.

Limits the bandwidth to reduce display

On

noise; filters the signal to reduce noise and

BW limit

other unwanted high frequency

Off

components.

Selects the resolution of the Volts/Div knob

Coarse         Coarse defines a 1-2-5 sequence.

Volts/Div

Fine           Fine changes the resolution to small steps

between the coarse settings.

1X,5X

Set to match the type of probe you are

10X,50X

Probe                         using to ensure correct vertical readouts.

100X,

500X,1000X

Next Page Page 1/3            Enter the second page of the menu.

Table 2-4 CH1, CH2 function menu 2:

Option          Setting       Instruction

on            Turn on invert function.

Invert

off           Turn off invert function.

Press this button to enter the “Digital Filter

Digital Filter

menu”.(See table 2-5)

Next Page Page 2/3           Enter the second page of the menu.

Table 2-5 Digital Filter function menu:

Option          Setting       Introduction

On       Turn on the digital filter.

Digital Filter

Off      Turn off the digital filter.

Setup as LPF (Low Pass Filter).

Setup as HPF (High Pass Filter).

Type

Setup as BPF (Band Pass Filter).

Setup as BRF (Band Reject Filter).

Upper_limit                  Turn the “Universal” knob to set upper limit.

Lower_limit                  Turn the “Universal” knob to set lower limit.

Return                     Return the digital filter main menu.

16 SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

Setting CH1, CH2 Channels

Each channel has its own separate Menu. The items are set up separately

according to each channel.

1. Choosing Coupling

Take the CH1 for example; the tested signal is a sine wave signal with DC

deflection:

● Press“CH1”→“Coupling”→“AC”, Set to AC couple mode. It blocks the DC

component of the input signal.

● Press“CH1”→“Coupling”→“DC”, Set to DC couple mode. Both DC and

AC component could be obstructed.

● Press“CH1”→“Coupling”→“GND”, Set to GROUND mode. It disconnects the

input signal.

Set to DC

Coupling

Picture 2.6-2

DC Status

2. Bandwidth Limiting

Take the CH1 for example; the tested signal is a pulse signal with the high

frequency surge:

● Press “CH1”→“BW Limit”→ “On”,Set the band width Limited to open state. The

high frequency component which is higher than 20MHz obstructed.

● Press“CH1”→“BW Limit”→ “Off”, Set bandwidth Limited to close state, the High

Frequency component in the tested signal could pass.

SDS1000CML+/SDS1000DL+    User Manual   17

SIGLENT

Set BW

to 20MHz

Picture 2.6-3

BW Limit Symbol

3. Adjust Sensitivity

Vertical scale adjusting has Coarse and Fine two modes, Vertical sensitivity range

is 2mV/div~10V/div scale.

Take the CH1 for example:

● Press “CH1”→“Volts/Div”→“Coarse”. It is the default setting of Volts/Div, and it

makes the vertical scaling in a 1-2-5-step sequence from 2mv/div, 5mv/div,

10mv/div to 10v/div.

● Press “CH1”→“Volts/Div”→“Fine”. This setting changes the vertical to small

steps between the coarse settings. It will be helpful when you need to adjust the

waveform vertical size in smooth steps.

Set to

Coarse

Picture 2.6-4

18  SDS1000CML+/SDS1000DL+      User Manual

SIGLENT

4. Setting Probe Attenuation

In order to assort the attenuation coefficient, you need to response in the channel

operation Menu. If the attenuation coefficient is 10:1, the input coefficient should

be set to 10X, so that the mistake of the Volts/div information and measure testing

should be forbidden.

Take the CH1 for example, when you use the 100:1 probe:

● Press“CH1”→“Probe” →“100”

Probe

Attenuation

Factor

Picture 2.6-5

Status of 50X

5. Inverting waveforms

Take the CH1 for example:

● Press“CH1”→Next Page“ page1/3” →“Invert”→“On”:

Picture 2.6-6

SDS1000CML+/SDS1000DL+       User Manual    19

SIGLENT

6. Using the Digital Filter

Press “CH1”→“Next Page page1/3”→ “Filter”, display the digital filter menu. Select

“Filter Type”, then select “Upper Limit” or “Lower Limit” and turn the “Universal”

knob to adjust them.

● Press “CH1”→“Next Page page1/3”→ “Filter” →“Off”. Turn off the Digital Filter

function.

Picture 2.6-7

● Press “CH1”→ “Next Page page1/3”→ “Filter” → “On”. Turn on the Digital

Filter function.

Picture 2.6-8

20   SDS1000CML+/SDS1000DL+     User Manual

SIGLENT

2.6.2 Using Vertical “Position” Knob and “Volt/div” Knob

■ Vertical “POSITION” Knob

1. Use the Vertical “POSITION” knobs to move the channel waveforms up or down

on the screen. This button’s resolution is variety as per the vertical scale.

2. When you adjust the vertical position of channels waveforms, the vertical

position information will display on the left bottom of screen. For example

“Volts Pos=24.6mV”.

3. Press the vertical “POSITION” knob to set the vertical position to zero.

■ “Volts/div” Knob

1. Use the “Volts/div” knobs to control how the oscilloscope amplifies or

attenuates the source signal of channel waveforms. When you turn the “volts/div”

knob, the oscilloscope increases or decreases the vertical size of the waveform on

the screen with respect to the ground level;

2. When you press the “Volt/div” Knob, you can switch “Volt/div” option between

“Coarse” and “Fine”. The vertical scale is made sure by the 1-2-5 step in the

Coarse. Increase in the clockwise, reduce in the anticlockwise. In the fine mode,

the knob changes the Volts/Div scale in small steps between the coarse settings.

Increase in the clockwise, reduce in the anticlockwise.

2.6.3 Math Functions

Math shows the results after +,-,*, / and FFT operations of the CH1 and CH2.

Press the MATH button to display the waveform math operations. Press the

MATH button again to remove the math waveform display.

Table 2-6    MATH function menu:

Function          Setting         Introduction

Math operates between signal source

Operation     +、-、*、/、FFT

CH1 and CH2.

on           Invert the MATH waveform.

Invert                          Turn off MATH Invert function.

off

Adjust the vertical position of the math

waveform by using universal knob.

Adjust the range of the math

waveform by using universal knob.

SDS1000CML+/SDS1000DL+        User Manual  21

SIGLENT

Table 2-7 MATH function instruction

Operation          Setting         Introduction

+            CH1+CH1          CH1 waveform adds CH2 waveform.

The channel 2 waveform is subtracted

CH1-CH2

from the channel 1 waveform.

-

The channel 1 waveform is subtracted

CH2-CH1

from the channel 2 waveform.

*            CH1*CH2          Channel 1 multiply Channel 2.

CH1/CH2          Channel 1 divides Channel 2.

/

CH2/CH1          Channel 2 divides Channel 1.

FFT      Fast Fourier Transform.

Picture 2.6-9 Math Waveform

1. FFT Spectrum Analyzer

The FFT process mathematically converts a time-domain signal into its frequency

components. You can use the Math FFT mode to view the following types of

signals:

● Analyze the Humorous wave in the Power cable.

● Test the Humorous content and distortion in the system

● Show the Noise in the DC Power supply

● Test the filter and pulse response in the system

● Analyze vibration

22  SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

Table 2-8 FFT function menu 1:

FFT Option       Setting         Introduction

Select this channel as the FFT source.

Source           CH1, CH2

Hanning         Select FFT window types.

Hamming

Window

Rectangular

Blackman

1X              Changes the horizontal magnification of

2X              the FFT display.

FFT ZOOM

5X

10X

Next Page        Page 1/2        Enter the second page of FFT menu.

Table 2-9 FFT function menu 2:

FFT Option        Setting      Introduction

Vrms         Set Vrms to be the Vertical Scale unit.

Scale

dBVrms       Set dBVrms to be the vertical Scale unit.

Split        Display FFT waveform on half screen.

Display

Full screen Display FFT waveform on full screen.

Next Page         Page 2/2     Return the first page of FFT menu.

To use the Math FFT mode, you need to perform the following tasks:

1. Set up the source (time-domain) waveform.

● Press the AUTO button to display a YT waveform.

● Turn the vertical “POSITION” knob to move the YT waveform to the center

vertically (zero divisions).

● Turn the horizontal “POSITION” knob to position the part of the YT waveform

that you want to analyze in the center eight divisions of the screen.

The oscilloscope calculates the FFT spectrum using the center 1024 points of

the time-domain waveform.

● Turn the “Volts/div” knob to ensure that the entire waveform remains on the

screen.

● Turn the “S/div” knob to provide the resolution you want in the FFT spectrum.

● If possible, set the oscilloscope to display many signal cycles.

To display FFT correctly, follow these steps:

1. Push the “MATH” button.

2. Set the “Operation” option to FFT.

SDS1000CML+/SDS1000DL+      User Manual  23

SIGLENT

3. Press “Source” button to select “CH1” or “CH2” according to input signal

channel.

4. According to Nyquist law, turn the “S/div” knob to adjust the sampling rate

(This parameter is displayed behind the time base parameter) is at least

double than input signal frequency.

2. Displaying the FFT Spectrum

Press the MATH button to display the Math Menu. Use the options to select the

Source channel, Window algorithm, and FFT Zoom factor. You can display only

one FFT spectrum at a time. You can select “Full screen” or “Split” in “Display”

option to display FFT waveform on full screen or display channel waveform and its

FFT waveform on half screen at a time.

Picture2.6-10 FFT function

3. Select FFT window

Windows reduce spectral leakage in the FFT spectrum. The FFT assumes that

the YT waveform repeats forever. With an integral number of cycles, the YT

waveform starts and ends at the me amplitude and there are no discontinuities in

the signal shape A non-integral number of cycles in the YT waveform causes the

signal start and end points to be at different amplitudes. The transitions between

the start and end points cause discontinuities in the signal that introduce

high-frequency transients.

According to the tested options and source speciality, make sure the window

you need to use.

24   SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

Table 2-10 FFT window instruction

Window                 Speciality             Satisfied Test content

Best frequency               Symmetric transients or

resolution, worst            bursts.

magnitude resolution.        Equal-amplitude sine waves

Rectangulr        This is essentially the      with fixed frequencies.

same as no window.           Broadband random noise with

a relatively slowly varying

spectrum.

Better frequency, poorer Sine, periodic, and

magnitude accuracy           narrow-band random noise.

Hanning          than Rectangular.            Asymmetic transients or

Hamming           Hamming has slightly         bursts.

better frequency

resolution than Hanning.

Best magnitude, worst        Single frequency waveforms,

Blackman

frequency resolution.        to find higher order harmonics.

4. Magnifying and Positioning an FFT Spectrum

You can magnify and use cursors to take measurements on the FFT spectrum.

The oscilloscope includes an “FFT Zoom” option to magnify horizontally, press

this option button to select “1X”, “2X”, “5X” or “10X”. Moreover, you also can turn

the “Universal” knob to magnify FFT waveform horizontally in a 1-2-5 step. To

magnify vertically; you can turn the “Volts/div” knob.

5. Measuring an FFT Spectrum Using Cursors

You can take two measurements on FFT spectrums: magnitude (in dB) and

frequency (in Hz). Magnitude is referenced to 0 dB, where 0 dB equals 1 VRMS.

You can use the cursors to take measurements at any zoom factor. (Refer to

cursor measure2.11.2)

Use horizontal cursors to measure amplitude and vertical cursors to measure

frequency.

SDS1000CML+/SDS1000DL+         User Manual   25

SIGLENT

If you input a sine signal to channel 1, follow these steps:

1. Measure FFT Amplitude

1) Input a sine signal to channel 1, and press the “AUTO” button.

2) Press the “MATH” button to enter the “MATH” menu.

3) Press the “Operation” option button to select “FFT”.

4) Press the “Source” option button to select “CH1”.

5) Press CH1 button to display CH1 menu.

6) Turn the “S/div” knob to adjust sampling rate (at least double bigger than

frequency of input signal).

7) If FFT display on full screen, press CH1 button again to remove channel

waveform display.

8) Press the “CURSOR” button to enter “Cursor” menu.

9) Press the “Cursor Mode” button to select “Manual”.

10) Press the “Type” option button to select “Voltage”.

11) Press the “Source” option button to select “MATH”.

12) Press the “CurA” option button; turn the “Universal” knob to move Cursor A to

the highest point of the FFT waveform.

13) Press the “CurB” option button, turn the “Universal” knob to move Cursor B to

the lowest point of the FFT waveform.

14) The amplitude (△T) displays on the top of the left screen.

Picture 2.6-11

26   SDS1000CML+/SDS1000DL+       User Manual

SIGLENT

2.   Measure FFT Frequency

1)   Press the CURSOR button.

2)   Press the “Cursor Mode” button to select “Manual”.

3)   Press the “Type” option button to select “Time”.

4)   Press the “Source” option button to select “MATH”.

5)   Press the “CurA” option button, turn the “Universal” button to move Cursor A

to the highest position of the FFT waveform.

6)   The value of CurA displaying on the top of the left screen is FFT highs

frequency. This frequency should be the same as input signal frequency.

Picture 2.6-12

2.6.4 Using Ref

The reference control saves waveforms to a nonvolatile waveform memory. The

reference function becomes available after a waveform has been saved.

Table 2-11 REF function menu:

Option Setting            Introduction

CH1,CH2,        Choose the waveform display to store.

Source      CH1 off

CH2 off

REFA                        Choose the reference location to store or

REFB                        recall a waveform.

Stores source waveform to the chosen

Save

reference location.

REFA        on              Recall the reference waveform on the screen.

REFB        off             Turn off the reference waveform.

SDS1000CML+/SDS1000DL+     User Manual 27

SIGLENT

Press the Ref button to display the “Reference waveform menu”.

Picture 2.6-13

Operation step:

1.   Press the “REF” menu button to display the “Reference waveform menu”.

2.   Press the “Source” option button to select input signal channel.

3.   Turn the vertical “POSITION” knob and “Volt/div” knob to adjust the

vertical position and scale to conformable positions.

4.  Press the third option button to select “REFA” or “REFB” as storage position.

5.  Press the “Save” option button.

6.  Press the bottom option button to select “REFA On” or “REFB On” to

recall the reference waveform.

Picture 2.6-14

28 SDS1000CML+/SDS1000DL+          User Manual

SIGLENT

2.7 Horizontal System

As follow Picture, there are one button and two knobs in the HORIZONTAL area.

S/div knob

Horizontal

position knob

Picture 2.7-1

Table 2- 12 SDS1000CML Horizontal system function menu:

Option        Setting    Description

Turn on this function that main timebase

On

waveform display on the top half screen

and window timebase waveform display on

Delayed

the below half screen at the same time.

Turn off this function that only display main

Off

timebase waveform on the screen.

Set memory depth to normal.

Normal

MemDepth                     Set memory depth to long Memory depth in

Long Mem

order to get more waveform dots.

Note: If Set memory depth to normal mode, the memory depth is 40Kpts

If Set memory depth to Long Mem mode, the memory depth is 2Mpts

MemDepth Detailed Specification:

Memory                           Samping      Short

Channel Mode                                   Long Memory

Depth                              Rate       Memory

Single Channel    1Gsa/s         40kpts        No Support

500MSa/s

Single Channel                   20kpts        2Mpts

or lower

Double            500MSa/s

20kpts        1Mpts

Channels          or lower

SDS1000CML+/SDS1000DL+        User Manual 29

SIGLENT

2.7.1 Horizontal Control Knob

You can use the horizontal controls to change the horizontal scale and position of

waveforms. The horizontal position readout shows the time represented by the

center of the screen, using the time of the trigger as zero. Changing the horizontal

scale causes the waveform to expand or contract around the screen center.

■ Horizontal “POSITION” Knob

1. Adjust the horizontal position of all channels and math waveforms (the position

of the trigger relative to the center of the screen). The resolution of this control

varies with the time base setting.

2. When you press the horizontal “POSITION” Knob, you can set the horizontal

position to zero.

■“S/div” Knob

1. Using to change the horizontal time scale to magnify or compress the waveform.

If waveform acquisition is stopped (using the RUN/STOP or SINGLE button), turn

the S/div knob to expand or compress the waveform.

2. Select the horizontal time/div (scale factor) for the main or the window time

base. When Window Zone is enabled, it changes the width of the window zone by

changing the window time base.

■ Display Scan Mode:

When the SEC/DIV control is set to100 ms/div or slower and the trigger mode is

set to Auto, the oscilloscope enters the scan acquisition mode. In this mode, the

waveform display updates from left to right. There is no trigger or horizontal

position control of waveforms during scan mode.

2.7.2 Window Zone

Use the Delayed option to define a segment of a waveform to see more detail. The

Window time base setting cannot be set slower than the Main time base setting.

You can turn the Horizontal Position and SEC/DIV controls to enlarge or minish

waveforms in the Window Zone.

“M” Mean main time base, “W” mean window time base.. It is also a arrow on the

scale top to show the vertical position.

30   SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

If you want to see a section of the waveform in detail, follow these steps:

(1) Press the “HORI MENU” button to enter the “Horizontal menu”.

(2) Turn the “S/div” knob to change the main timebase scale.

(3) Press the “Delayed” option button to select “On”.

Picture2.7-2

(4) Turn the “Horizontal Position” knob (adjust window’s position) to select

the window that your need and expanded window waveform display on

the below half screen at the same time.

SDS1000CML+/SDS1000DL+       User Manual   31

SIGLENT

2.8 Trigger System

The trigger determines when the oscilloscope starts to acquire data and display a

waveform. When a trigger is set up properly, the oscilloscope converts unstable

displays or blank screens into meaningful waveforms.

There are three buttons and one knob in the Trigger area. See picture.2-29:

LEVEL knob

Picture 2.8-1

■“TRIG MENU” Button: Press the “TRIG MENU” button to display “Trigger

Menu”.

■“LEVEL” Knob: The LEVEL knob is to set the corresponding signal voltage of

trigger point in order to sample. Press the “LEVEL” knob can set trigger level to

zero.

■“SET TO 50%” Button: Use the “SET TO 50%” button to stabilize a waveform

quickly. The oscilloscope can set the Trigger Level to be about halfway between

the minimum and maximum voltage levels automatically. This is useful when you

connect a signal to the EXT TRIG BNC and set the trigger source to Ext or Ext/5.

■“FORCE” Button: Use the FORCE button to complete the current waveform

acquisition whether the oscilloscope detects a trigger or not. This is useful for

SINGLE acquisitions and Normal trigger mode.

■ Pre-trigger/Delayed trigger: The data before and after trigger the trigger

position is typically set at the horizontal center of the screen, in the full-screen

display the 6div data of pre-trigger and delayed trigger can be surveyed. More

data of pre-trigger and 1s delayed trigger can be surveyed by adjusting the

horizontal position.

32   SDS1000CML+/SDS1000DL+       User Manual

SIGLENT

The feature is very useful because you can see the events that led up the trigger

point everything to the right of the trigger point is called posttrigger information the

amount of delay range (pre-trigger and posttrigger information) available is

dependent on the sweep speed selected.

2.8.1 Signal Source

You can use the Trigger Source options to select the signal that the oscilloscope

uses as a trigger. The source can be any signal connected to a channel BNC, to

the EXT TRIG BNC or the AC power line (available only with Edge triggers).

SDS1000CML+/SDS1000DL+         User Manual   33

SIGLENT

2.8.2 Trigger Type

The scopes have five trigger types: Edge, Video, Pulse, Slope, and

Alternative.

■ Edge Trigger

Use Edge triggering to trigger on the edge of the oscilloscope input signal at the

trigger threshold.

Table 2-13 Edge Trigger function Menu:

Option Setting Explain

With Edge highlighted, the rising or falling edge of

Type     Edge

the input signal is used for the trigger.

CH1     Triggers on a channel whether or not the waveform

CH2     is displayed.

Does not display the trigger signal; the Ext option

uses the signal connected to the EXT TRIG

EXT

front-panel BNC and allows a trigger level range of

Source             -1.2V to +1.2V.

Same as Ext option, but attenuates the signal by a

EXT/5 factor of five, and allows a trigger level range of +6V

to -6V.This extends the trigger level range.

This selection uses a signal derived from the power

AC Line line as the trigger source; trigger coupling is set to

DC and the trigger level to 0 volts.

Trigger on Rising edge of the trigger signal.

Trigger on Falling edge of the trigger signal.

Slope

Trigger on Rising edge and Falling edge of the

trigger signal.

Use this mode to let the acquisition free-run in the

absence of a valid trigger; This mode allows an

Auto

untriggered, scanning waveform at 100 ms/div or

slower time base settings.

Use this mode when you want to see only valid

Mode

triggered waveforms; when you use this mode, the

Normal

oscilloscope does not display a waveform until after

the first trigger.

When you want the oscilloscope to acquire a single

Single

waveform, press the “SINGLE ”button.

Set up             Enter the “Trigger Setup Menu” (See table 2-14).

34   SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

Table 2-14 Trigger Setup function menu

Option      Setting      Explain

Coupling        DC        Passes all components of the signal

Blocks DC components ,attenuates signals

AC

below 50 Hz.

Attenuates the high-frequency components

HF Reject

above 150 kHz.

Blocks the DC component , attenuates the

LF Reject

low-frequency components below 7 kHz.

Holdoff                  Using the “universal” knob to adjust holdoff

time(sec),the holdoff value is displayed.

Holdoff                    Reset holdoff time to 100ns.

Reset

Return                   Return the first page of “Trigger main menu”.

Picture2.8-2

Operate Instruction:

1. Set up Type

1) Press the “TRIG MENU” button to display “Trigger” menu.

2) Press the “Type” option button to select “Edge”.

2. Set up Source

According to input signal, press the “Source” option button to select

“CH1”, “CH2”,“EXT”, “EXT/5”or “AC Line”.

3. Set up Slope

Press the “Slope” option button to select “        ”, “     ” or “       ”.

4. Set up Trigger mode

Press the “Trigger mode” option button to select “Auto”, “Normal”,

“Single”.

SDS1000CML+/SDS1000DL+       User Manual    35

SIGLENT

Auto: The waveform refresh at a high speed whether the trigger

condition is satisfied or not.

Normal: The waveform refresh when the trigger condition is satisfied

and waits for next trigger event occurring when the trigger condition is

not satisfied.

Single: The oscilloscope acquire a waveform when the trigger

condition is satisfied and then stops.

5. Set up Trigger coupling

a. Press the “Set up” button to enter the “Trigger Setup Menu”.

b. Press the “Coupling” option button to select “DC”, “AC”, “HF

Reject” or “LF Reject”.

36 SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

■ Pulse Trigger:Use Pulse Width triggering to trigger on aberrant

pulses.

Table 2-15   Pulse Trigger function Menu 1:

Option                 Setting              Explain

Select the pulse to trigger

Type                   Pulse               the pulse match the

trigger condition.

CH1                Select input signal source.

CH2

Source

EXT

EXT/5

(Positive pulse width less   Select how to compare

than pulse width setting)         the trigger pulse relative to

(Positive pulse width larger the value selected in the

than pulse width setting)         Set Pulse Width option.

(Positive pulse width equal

to pulse width setting)

When

(Negative pulse width less

than pulse width setting)

(Negative pulse width

larger than pulse width setting)

(Negative pulse width

equal to pulse width setting)

Selecting this option can

Set

20.0ns~10.0s             turn the universal to set up

Width

the pulse width.

Next                                       Press this button to enter

Page 1/2

Page                                       the second page.

Picture 2.8-3

SDS1000CML+/SDS1000DL+      User Manual   37

SIGLENT

Table 2-16   Pulse Trigger function Menu 2:

Option        Setting     Explain

Select the pulse to trigger the pulse match

Type          Pulse

the trigger condition.

Auto      Select the type of triggering; Normal mode is

Mode           Normal     best for most Pulse Width trigger

single     applications.

Set up                    Enter the “Trigger setup menu”.

Next Page       Page 2/2    Press this button to return the first page.

Picture 2.8-4

Operate Instruction:

1. Set up Type

1) Press the “TRIG MENU” button to display “Trigger menu”.

2) Press the “Type” option button to select “Pulse”.

2. Set up condition

Press the “When” option button to select “     ”、“      ”、“     ”、“ ”、

“   ”or“    ”。

3. Set up pulse width

Turn the “Universal” knob to set up width.

38 SDS1000CML+/SDS1000DL+       User Manual

SIGLENT

■ Video Trigger:Trigger on fields or lines of standard video signals.

Table 2-17 Functional Manu of Video Trigger 1:

Option       Setting          Instruction

When you select the video type, put the

couple set to the AC, then you could

Type         Video

trigger the NTSC,PAL and SECAM video

signal.

CH1              Select the input source to be the trigger

CH2              signal.

Source

EXT              Ext and Ext/5 use the signal applied to

EXT/5            the EXT TRIG connector as the source.

(Normal)   Normal triggers on the negative edge of

the sync pulse.

Polarity

(Inverted) Inverted triggers on the positive edge of

the sync pulse.

Line Num

Sync         All lines        Select appropriate video sync.

Odd field

Even Field

Enter the second page of “Video

Next Page    Page 1/2

trigger menu”.

Table 2-18 Functional Manu of Video Trigger 2:

Option       Setting            Instruction

When you select the video type, put

the couple set to the AC, then you

Type         Video

could trigger the NTSC,PAL and

SECAM video signal.

NTSC               Select the video standard for sync

Standard

Pal/Secam          and line number count.

Use this mode to let the acquisition

free-run in the absence of a valid

trigger; This mode allows an

Auto

untriggered, scanning waveform at

100 ms/div or slower time base

settings.

Use this mode when you want to see

Mode

only valid triggered waveforms; when

Normal             you use this mode, the oscilloscope

does not display a waveform until

after the first trigger.

when you want the oscilloscope to

Single             acquire a single waveform, press the

“SINGLE ”button.

Set up                          Enter the “Trigger setup menu”.

Return the first page of “Video

Next Page    Page 2/2

Trigger menu”.

SDS1000CML+/SDS1000DL+      User Manual 39

SIGLENT

Picture 2.8-5

Operate Instruction

1. Set up Type

1) Press the “TRIG MENU” button to display “Trigger” menu.

2) Press the “Type” option button to select “Video”.

2. Set up Polarity

Press the “Polarity” option button to select “   ” or “    ”.

3. Set up synchronization

1) Press the “Sync” option button to select “All Lines”, “Line Num”, “Odd

Field”, and “Even Field”.

2) If you select “Line Num”, you can turn the “Universal” knob to set the

appointed line number.

4. Set up Standard

1) Press the “Next Page Page 2/2” option button.

2) Press the “Standard” option button to select “PAL/SECAM” or “NTSC”.

40 SDS1000CML+/SDS1000DL+       User Manual

SIGLENT

■ Slope Trigger:Trigger on positive slope of negative slope according

to setup time of the oscilloscope.

Table 2-19 Slope trigger function menu 1

Option         Setting         Instruction

Trigger on positive slope of negative

Type           Slope           slope according to setup time of the

oscilloscope.

CH1             Select trigger source.

CH2

Source

EXT

EXT/5

Select trigger condition.

When

Turn the “Universal” knob to set slope

Time

〈Set time〉      time. Time setup range is 20ns-10s.

Next Page      Page 1/2        Enter the second page of slope trigger.

Picture 2.8-6

SDS1000CML+/SDS1000DL+     User Manual    41

SIGLENT

Table 2-20 Slope trigger function menu 2

Option       Setting     Instruction

Trigger on positive slope of negative slope

Type         Slope

according to setup time of the oscilloscope.

Vertical                 Select the trigger level that can be adjusted by

“LEVEL” knob. You can adjust “LEVEL A”,

“LEVEL B” or adjust them at the same time.

Use this mode to let the acquisition free-run in

the absence of a valid trigger; This mode

Auto

allows an untriggered, scanning waveform at

100 ms/div or slower time base settings.

Mode

Use this mode when you want to see only valid

triggered waveforms; when you use this mode,

Normal

the oscilloscope does not display a waveform

until after the first trigger.

When you want the oscilloscope to acquire a

Single

single waveform, press the “SINGLE” button.

Set up                   Enter “Trigger setup menu” (See table 2-14).

Next Page    Page 2/2    Return the first page of slope trigger.

Picture 2.8-7

Operate Instruction:

Follow the next steps if you select “Slope trigger”:

1.   Input a signal to CH1 or CH2.

2.   Press the “AUTO” button.

3.   Press the “TRIG MENU” button to enter “Trigger menu”.

4.   Press the “Type” option button to select “Slope”.

5.   Press the “Source” option button to select “CH1” or “CH2”.

42  SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

6.  Press the “When” option button to select “         ”, “       ”, “       ”,

“       ”,“         ”或 “            ”.

7.  Press the “Time” button, turn the “Universal” knob to adjust slope time.

8.  Press the “Next Page     Page 1/2” option button to enter the second page

of “Slope trigger menu”.

9.  Press the “Vertical” option button to select trigger level that can be adjusted.

10. Turn the “LEVEL” knob.

SDS1000CML+/SDS1000DL+            User Manual   43

SIGLENT

■ Alternative trigger

The trigger signal comes from two vertical channels when you use alternative

trigger. In this mode, you can observe two irrelative signals at the same time. You

can select different trigger types for two vertical signals, and selected types cover

edge, pulse, video and slope trigger. Trigger information of two channel signals

display on the bottom right of the screen.

Picture 2.8-8

Table 2-21 Set trigger mode to edge trigger function menu 1:

Option Setting          Instruction

When use alternative trigger, the trigger signal

Type      Alternative comes from two vertical channels. In the mode,

you can observe two irrelative signals at a time.

Chann

CH1-CH2 Set the trigger channels

els

CH1       Set trigger type information for CH1 signal

Source

CH2       Set trigger type information for CH2 signal

Mode         Edge      Set trigger type of vertical channel signal to edge

Triggering on rising edge.

Slope

Triggering on falling edge.

Triggering on rising edge and falling edge.

Set up                  Enter “Trigger setup menu” (See table 2-14).

44   SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

Table 2-22 Set trigger mode to pulse trigger function menu 1:

Option      Setting        Instruction

The trigger signal comes from two vertical

Alternative channels when you use alternative trigger. In

Type

this mode, you can observe two irrelative

signals at the same time.

CH1            Set trigger type information for CH1 signal

Source

CH2            Set trigger type information for CH2 signal

Set trigger type of the vertical channel signal

Mode        Pulse

to Pulse trigger.

Select how to compare the trigger pulse

When                       relative to the value selected in the Set Pulse

Width option.

Enter the second page of Alternative trigger

Next Page   Page 1/2

menu.

Table 2-23 Set trigger mode to pulse trigger function menu 2:

Option      Setting          Explain

Set Width   20.0ns       ~ Selecting this option can turn the universal

10.0s            to set up the pulse width.

Set up                       Enter the “Trigger Setup Menu”(see table

2-14).

Next Page Page 2/2           Press this button to return the first page.

Table 2-24 Set trigger mode to video trigger function menu1:

Option     Setting            Instruction

Type       Alternative        The trigger signal comes from two vertical

channels when you use alternative trigger.

In this mode, you can observe two irrelative

signals at the same time.

CH1                Set trigger type information for CH1 signal

Source

CH2                Set trigger type information for CH2 signal

Set trigger type of the vertical channel

Mode       Video

signal to Video trigger.

Normal triggers on the negative edge of the

(Normal)     sync pulse.

Polarity                      Inverted triggers on the positive edge of the

(Inverted) sync pulse.

Next                          Enter the second page of Alternative

Page 1/2

Page                          trigger menu.

SDS1000CML+/SDS1000DL+        User Manual  45

SIGLENT

Table 2-25 Set trigger mode to video trigger function menu 2:

Option    Setting          Instruction

Sync      Line Num

All lines        Select appropriate video sync.

Odd field

Even Field

Standard NTSC              Select the video standard for sync and line

Pal/Secam        number count.

Next                       Enter the second page of Alternative trigger

Page 1/2

Page                       menu.

Table 2-26 Set trigger mode to slope trigger function menu 1

Option     Setting      Instruction

Type       Alternative The trigger signal comes from two vertical

channels when you use alternative trigger. In

this mode, you can observe two irrelative

signals at the same time.

Set trigger type informa

CH1

Source                  tion for CH1 signal

CH2

Set trigger type information for CH2signal

Set trigger type of the vertical channel signal

Mode       Slope

to slope trigger.

When                    Select slope trigger condition.

Enter the second page of the alternative

Next Page Page 1/2

trigger.

Table 2-27 Set trigger mode to slope trigger function menu 2

Option       Setting      Instruction

Time                      Turn the “Universal” knob to set the slope

〈       Set time. Time setup range is 20ns-10s.

time〉

Vertical                  Select the trigger level that can be adjusted

by “LEVEL” knob. You can adjust “LEVEL

A”, “LEVEL B” or adjust them at the same

time.

Set up                    Enter “Trigger setup menu”(see table 2-14).

Next Page    Page 2/2     Return to the first page of “Alternative trigger

menu”.

46  SDS1000CML+/SDS1000DL+     User Manual

SIGLENT

Operate Instruction:

Observe two irrelative channel signals, follow these steps:

1. Input two irrelative signals to channel 1and channel 2.

2. Press the AUTO button.

3. Press the TRIG MENU button to enter “trigger menu”.

4. Press the “Type” option button to select “Alternative”.

5. Press the “Channels” option button to select “CH1-CH2”

6. Press the “Source” option button to select “CH1”.

7. Press the CH1 button and turn the “S/div” knob to optimize waveform

display.

8. Press “Mode” option button to select “Edge”, “Pulse”, “Slope” or “Video”.

9. Set the trigger according to trigger edge.

10. Press the “Source” option button to select “CH2”.

11. Press the CH2 button and turn the “S/div” knob to optimize waveform

display.

12. Repeat steps 8 and 9.

2.8.3 Coupling

Use the “Coupling” to make sure the signal that passes the trigger circuit. It is

useful for us to gather a steady wave form.

If you use the trigger coupling, you should press the “TRIGGER” button and then

select “edge”, “Pulse”, “Video” or “Slope” trigger. Then select the “coupling” option

in “Set up menu”.

2.8.4 Position

The horizontal position control establishes the time between the trigger position

and the screen center. You can adjust the horizontal “POSITION” knob control to

view waveform data before the trigger, after the trigger, or some of each. When

you change the horizontal position of a waveform, you are changing the time

between the trigger and the center of the display actually. (This appears to move

the waveform to the right or left on the display.)

SDS1000CML+/SDS1000DL+        User Manual   47

SIGLENT

2.8.5 Slope & Level

The Slope and Level controls help to define the trigger. The Slope option (Edge

trigger type only) determines whether the oscilloscope finds the trigger point on

the rising or the falling edge of a signal.

The TRIGGER LEVEL knob controls where on the edge the trigger point occurs.

Trig ge r le ve l ca n be ad ju ste d vert ica lly

Rising edge              Falling edge

Picture 2.8-9

Note: ● Press the SINGLE button when you want the oscilloscope to

acquire a single waveform.

● Trigger coupling affects only the signal passed to the trigger

system. It does not affect the bandwidth or coupling of the

signal displayed on the screen.

● Normal Polarity Sync triggers always occur on

negative-going horizontal sync pulses. If the video

waveform has positive-going horizontal sync pulses, use

the Inverted Polarity selection.

48   SDS1000CML+/SDS1000DL+                                                  User Manual

SIGLENT

2.8.6 Trigger Holdoff

You can use the Trigger Holdoff function to produce a stable display of complex

waveforms. Holdoff is time between when the oscilloscope detects one trigger and

when it is ready to detect another. The oscilloscope will not trigger during the

holdoff time. For a pulse train, you can adjust the holdoff time so the oscilloscope

triggers only on the first pulse in the train.

Trigger position

Trigger level

Holdoff time

Picture 2.8-10

If you want to change holdoff time, please follow next operations:

1. Press the “TRIG MENU” button to show the “TRIG Menu”.

2. Press the “Type” option button to select trigger type.

3. Press the “Set up” option button to enter the “Trigger setup menu”.

4. Press the “Holdoff” option button,turn the “Universal” knob to change the

holdoff time until the waveform trigger steadily.

Note:Use trigger holdoff to help stabilize the display of aperiodic waveforms.

SDS1000CML+/SDS1000DL+    User Manual   49

SIGLENT

2.9 Acquiring Signals System

Showing as the follow picture,the “ACQUIRE” button for Acquiring Signals system

is at the menu.

Picture 2.9-1

Table 2-28     The Function manual of Acquiring Signals:

Option        Setting      Introduction

Use for sampling and accurately display most

Sampling

of the waveform.

Peak      Detect the noise and decrease the

Acquisition

Detect     possibility of aliasing.

Use to reduce random or uncorrelated noise

Average

in the signal display.

Averages

〔4, 16,32

,64,128, Select number of averages.

256〕

On       Use sin interpolation

Sinx/x

off      Use linear interpolation

Equ time Set the Sampling mode to Equ time.

Mode

Real time Set the Sampling mode to Real time.

Sa Rate                  Display system sampling rate.

When you acquire a signal, the oscilloscope converts it into a digital form and

displays a waveform. The acquisition mode defines how the signal is digitized and

the time base setting affects the time span and level of detail in the acquisition.

■ Sampling: In this acquisition mode, the oscilloscope samples the signal in

evenly spaced intervals to construct the waveform. This mode accurately

represents signals most of the time.

Advantage: You can use this mode to reduce random noise.

Disadvantage: This mode does not acquire rapid variations in the signal that may

50   SDS1000CML+/SDS1000DL+       User Manual

SIGLENT

occur between samples. This can result in aliasing may cause narrow pulses to be

missed. In these cases, you should use the Peak Detect mode to acquire data.

Picture 2.9-2 Sampling modes

■ Peak Detect: Peak Detect mode capture the maximum and minimum values of

a signal Finds highest and lowest record points over many acquisitions.

Advantage: In this way, the oscilloscope can acquire and display narrow pulses,

which may have otherwise been missed in Sample mode.

Disadvantage: Noise will appear to be higher in this mode.

Picture 2.9-3 Peak Detect mode

SDS1000CML+/SDS1000DL+    User Manual  51

SIGLENT

■  Average: The oscilloscope acquires several waveforms, averages them, and

displays the resulting waveform.

Advantage: You can use this mode to reduce random noise.

Picture 2.9-4 Average Mode

■ Equivalent Time Sampling:

The equivalent time sampling mode can achieve up to 20 ps of horizontal

resolution (equivalent to 50GSa/s). This mode is good for observing repetitive

waveforms.

■ Real Time Sampling: The scope has the highest Real-time sampling rate

up to 1GSa/s.

■ “RUN/STOP” Button: Press the RUN/STOP button when you want the

oscilloscope to acquire waveforms continuously. Press the button again to stop

the acquisition.

■ “SINGLE” Button: Press the SINGLE button to acquire a single waveform.

Each time you push the SINGLE button, the oscilloscope begins to acquire

another waveform. After the oscilloscope detects a trigger it completes the

acquisition and stops.

52  SDS1000CML+/SDS1000DL+       User Manual

SIGLENT

When you push the RUN/STOP or SINGLE buttons to start an acquisition, the

oscilloscope goes through the following steps:

1). Acquire enough data to fill the portion of the waveform record to the left of

the trigger point. This is also called the pre-trigger.

2). Continue to acquire data while waiting for the trigger condition to occur.

3). Detect the trigger condition.

4). Continue to acquire data until the waveform record is full.

5). Display the newly-acquired waveform.

■ Time Base: The oscilloscope digitizes waveforms by acquiring the value of an

input signal at discrete points. The time base allows you to control how often the

values are digitized.

Use S/div knob to adjust the time base to a horizontal scale suits your purpose,

■ Time Domain Aliasing:

Aliasing occurs when the oscilloscope does not sample the signal fast enough to

construct an accurate waveform record. When this happens, the oscilloscope

displays a waveform with a frequency lower than the actual input waveform, or

triggers and displays an unstable waveform.

Actual high-frequency waveform

Apparent low-frequency

waveform due to aliasing

Sampled points

Picture 2.9-5

SDS1000CML+/SDS1000DL+   User Manual    53

SIGLENT

Operate Introduction:

Set up Sampling Format

You can press the “Acquisition” option button or turn the “Universal” knob to select

“Sampling” mode, “Peak Detect” mode or “Average” mode..

Set up Averages

When you select “Average” format, you can press the “Averages” option button to

select “4”, “16”, “32”, “64”, “128”or “256”.

Set up function interpolation

You can also turn Sins/x interpolation on or off. Press the corresponding option

button to turn Sins/s on or off.

“Sinx” is sine interpolation

Picture 2.9-6

“X” is linear interpolation.

Picture 2.9-7

54  SDS1000CML+/SDS1000DL+          User Manual

SIGLENT

Set up Sampling Mode

Press the “Mode” option button to select “Real Time” or “Equ Time”.

Set up Sampling Rate

Adjust the sampling rate by pressing the “Sa Rate” option button and turning the

Time/div front panel knob. The sampling rate is shown at the corresponding

timebase scale.

SDS1000CML+/SDS1000DL+      User Manual   55

SIGLENT

2.10 Display System

The display function could be expressed by the “DISPLAY” Button.

Picture 2.10-1

Table 2-29 Display system function menu 1:

Option          Setting         Introduction

Vectors fill the space between adjacent

Vectors

sample points in the display.

Type

There is no link between adjacent

Dots

sample points.

Off

1 sec

Persist                         Sets the length of time each displayed

2 sec

sample point remains displayed.

5 sec

Infinite

Intensity                       Set waveforms’ intensity.

<Intensity>

Brightness                      Set grid brightness.

<Brightness>

Next Page       Page 1/3        Press this button to enter second page.

Picture 2.10-2

56 SDS1000CML+/SDS1000DL+       User Manual

SIGLENT

Table 2-30 Display system function menu 2:

Option         Setting                   Introduction

YT format displays the vertical voltage

in relation to time (horizontal scale). XY

Format         YT

format displays a dot each time a

XY

sample is acquired on channel 1 and

channel 2

Normal    Set to normal mode.

Screen

Inverted Set to invert color display mode.

Display grids and axes on the screen.

Grid                     Turn off the grids.

Turn off the grids and axes.

2sec

5sec

Menu Display         10sec    Set menu display time on screen.

20sec

Infinite

Press this button to enter the second

Next Page       Page 2/3

page of “Display menu”.

Picture 2.10-3

Table 2-31 Display system function menu3:

Option           Setting                      Introduction

Skin           Classical      Set up screen style.

Modern

Tradition

Succinct

Next Page         Page 3/3       Press this button to return the first page.

SDS1000CML+/SDS1000DL+         User Manual   57

SIGLENT

Operate Introduction:

1. Set up waveform display type

1) Press the “DISPLAY” button to enter the “Display” menu.

2) Press the “Type” option button to select “Vectors” or “Dots”.

2. Set up Persist

Press “Persist” option button to select “Off”, “1 Sec”, “2 Sec”, “5Sec” or

“Infinite”. You can use this option to observe some especial waveforms.

Picture 2.10-4

3. Set up Intensity

1) Press the “Intensity” option button and turn the “Universal” knob to adjust

waveforms’ intensity.

4. Set up Brightness

1) Press the “Brightness” option button and turn the “Universal” knob to

adjust grid brightness.

5. Set up display format

1) Press the “Next Page” option button to enter second display menu.

2) Press the “Format” option button to select “YT” or “XY”.

6. Set up Screen

Press the “Screen” option button to select “Normal” or “Inverted” to set the

screen display color.

7. Set up Grid

Press the “Grid” option button to select “         ”, “    ”or“      ”to set the

screen whether display grid or not.

58   SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

8. Set up Menu Display

Press the “Menu Display” option button to select “2 sec”, “5sec”, “10sec”,

“20sec” or “Infinite” to set menu display time on screen.

9. Set Skin

Press the “skin” option button or turn the “Universal” knob to select “Classical”,

“Modern”, “Traditional” or “Succinct”.

2.10.1 X-Y Format

Use the XY format to analyze phase differences, such as those represented by

Lissajous patterns. The format plots the voltage on channel 1 against the voltage

on channel 2, where channel 1 is the horizontal axis and channel 2 is the vertical

axis. The oscilloscope uses the untriggered Sample acquisition mode and

displays data as dots.

XY mode has a breakthrough that traditional oscilloscopes restrict sampling rate

at 1MSa/s,

support 25KSa/s~250MSa/s (1-2.5-5 sequence) adjusted ;

Operation steps :

●Channel 1 “Volt/div” and vertical “POSITION” set up the horizontal scale and

position.

●Channel 2 “Volt/div” and vertical “POSITION” set up the horizontal scale and

position.

●Turn the “S/div” knob to adjust the sampling rate

●The following functions are forbidden at XY display form:

●Benchmark wave form and wave mathematic

●Cursor

●Auto(resets display format to YT)

●Trigger Control

●Horizontal Position Knob

●Vector Display Type

●Scan Display

SDS1000CML+/SDS1000DL+      User Manual    59

SIGLENT

2.11 Measure System

The Oscilloscope displays the voltage in relation to time and test the wave form

displayed. There are scale, Cursor and auto measure modes.

2.11.1 Scale Measurement

This method allows you to make a quick, visual estimate. For example, you might

look at waveform amplitude and determine that it is a little more than 100 mV. You

can take simple measurements by counting the major and minor graticule

divisions involved and multiplying by the scale factor. For example, if you counted

five major vertical graticule divisions between the minimum and maximum values

of a waveform and knew you had a scale factor of 100mV/div, then you could

easily calculate your peak-to-peak voltage as follows:

5 divisions x 100 mV/division = 500 mV.

2.11.2 Cursor Measurement

Picture 2-50 displays the cursor button on the front-panel for this menu.

Picture 2.11-1

Press the “CURSORS” button to display the “Cursor” Menu.

The cursor measurement has three modes: Manual, Track and Auto Measure.

60   SDS1000CML+/SDS1000DL+         User Manual

SIGLENT

Manual Mode

Table 2-32 Manual cursor function menu:

Option         Setting     Instruction

Cursor Mode Manual In this menu, set the manual cursor measure.

Voltage Use cursor to measure voltage parameters.

Type

Time     Use cursor to measure time parameters.

CH1

CH2

Source        MATH      Select input signal channel.

REFA

REFB

Cur A                  Select this option,use “Universal” knob to

adjust cursor A.

Cur B                   Select this option,use “Universal” knob to

adjust cursor B.

In this mode, the screen displays two horizontal parallel cursors or vertical parallel

cursors to measure voltage or time. You can move the cursor by turning the

“Universal” knob. Before using cursors, you should make sure that you have set

the signal source as the channel for measuring.

■ Voltage Cursor: Voltage cursors appear as horizontal lines on the display and

measure the vertical parameters.

■ Time Cursor: Time cursors appear as vertical lines on the display and measure

the horizontal parameters.

■ Cursor Moving: Use the “universal” knob to move cursor1 and cursor2 .They

could be moved when the corresponding cursor option are selected, and cursor

value will display on the bottom left and top left of the screen when you move the

cursor.

To do manual cursor measurements, follow these steps:

1. Press CURSOR button to enter the cursor function menu.

2. Press the “Cursor Mode” option button to select “Manual”.

3. Press the “Type” option button to select “Voltage” or “Time”.

4. Press the “Source” option button to select “CH1”, “CH2”, “MATH”,

“REFA” ,“REFB”    according to input signal channel.

5. Select “Cur A”, turn the “Universal” knob to adjust Cursor A.

6. Select “Cur B”, turn the “Universal” knob to adjust Cursor B.

SDS1000CML+/SDS1000DL+       User Manual   61

SIGLENT

7. The measurement values are displayed on the top of the left corner:

If the measurement type is set to “Voltage”, the values are:

The value of Cur A: CurA

The value of Cur B: Cur B

The voltage increment between Cursor A and Cursor B: △V

If the measurement type is set to “Time”, the values are:

The value of Cur A: CurA

The value of Cur B: Cur B

The time increment between Cursor A and Cursor B: △T

The reciprocal of time increment between Cursor A and Cursor B:

1/△T

Picture 2.11-2

62 SDS1000CML+/SDS1000DL+      User Manual

SIGLENT

Track mode

Table 2-33 Track mode function menu:

Option           Setting     Instruction

Cursor Mode        Track    In this mode, set track cursor measure.

CH1      Set the input signal channel that the Cursor

Cursor A         CH2      A will measure.

NONE

CH1      Set the input signal channel that the Cursor

Cursor B         CH2      B will measure.

NONE

Cur A                    Select this option, turn the “Universal” knob

to adjust horizontal coordinate of Cursor A.

Cur B                    Select this option, turn the “Universal” knob

to adjust horizontal coordinate of Cursor B.

In this mode, the screen displays two cross cursors. The cross cursor sets the

position on the waveform automatically. You could adjust cursor’s horizontal

position on the waveform by turning the “Universal” knob. The oscilloscope

displays the values on the top of the right screen.

To do track cursor measurement, follow these steps:

1. Press CURSOR button to enter the cursor measure function menu.

2. Press the “Cursor Mode” option button to select “Track”.

3. Press the “Cursor A” option button to select the input signal channel.

4. Press the “Cursor B” option button to select the input signal channel.

5. Select “Cur A”, turn the “Universal” knob to move Cursor A

horizontally.

6. Select “Cur B”, turn the “Universal” knob to move Cursor B

horizontally.

7. The measurement values are displayed on the left of the top corner:

A→T: The horizontal position of Cursor A (Time cursor centered

around the midpoint of screen).

A→V: The Vertical position of Cursor A (Voltage cursor centered

around channel ground level).

B→T: The horizontal position of Cursor B (Time cursor centered

around the midpoint of screen).

SDS1000CML+/SDS1000DL+       User Manual   63

SIGLENT

B→V: The Vertical position of Cursor B (Voltage cursor centered

around channel ground level).

△T: Horizontal space between Cursor A and Cursor B (Time value

between two cursors).

1/△T: The reciprocal of horizontal space between cursor A and cursor

B.

△V: Vertical space between Cursor A and Cursor B (Voltage value

between two cursors).

Picture 2.11-3

64 SDS1000CML+/SDS1000DL+    User Manual

SIGLENT

Auto mode

Table 2-34 Auto mode function menu:

Option          Setting       Instruction

Cursor Mode          Auto        Set to auto cursor measure mode.

This mode will take effect with automatic measurements. The instruments will

display cursors while measuring parameters automatically. These cursors

demonstrate the physical meanings of these measurements.

To do auto cursor measurements, follow these steps:

1. Press the CURSOR button to enter “Cursor measure menu”.

2. Press the “Cursor Mode” option button to select “Auto”.

3. Press the “MEASURE” button to enter “Auto cursor measure mode

menu” to select the parameter that you want to measure.

Picture 2.11-4

SDS1000CML+/SDS1000DL+      User Manual 65

SIGLENT

2.11.3 Auto Measurement

Picture 2.11-5

When you take automatic measurements, the oscilloscope does all the calculating

for you. Because the measurements use the waveform record points, they are

more accurate than the graticule or cursor measurements.

Press the ‘MEASURE’ for the Automatic Test.

There are three auto measurement types: Voltage Measure, Time Measure and

Delay Measure. There are thirty two measure parameters in all.

Table 2-35 Auto measure function menu 1:

Option     Instruction

Voltage    Press this button to enter the Voltage measure menu.

Time       Press this button to enter the Time measure menu.

Delay      Press this button to enter the Delay measure menu.

All Mea    Press this button to enter the All Measurement menu.

Clear      Press this button to clear all the measurement that you

Measure    chose.

Picture 2.11-6

66   SDS1000CML+/SDS1000DL+     User Manual

SIGLENT

Table 2-36 Auto measure function 2-Voltage measure menu:

Option             Setting              Instruction

Select input signal source for

Source            CH1, CH2

Voltage measure.

Vmax, Vmin, Vpp, Vamp,       Press the “Type” button or

Vtop, Vbase, Cycle Mean,     turn the “Universal” knob to

Type      Mean, Cycle Vrms, Vrms,      select      Voltage     measure

ROVShoot, FOVShoot,         parameter.

RPREShoot, FPREShoot

Choose the measurement,

after pressing the button, the

Add                                    measurement value will be

shown         in    the    lower

left of the screen

Return to the first page of auto

Return

measurement menu.

Table 2-37 Auto measure function3-Time measure menu:

Option             Setting             Instruction

Select input signal source for

Source CH1, CH2

Time measure.

Rise Time、Fall Time、Freq、 Press the “Type” button or turn

Type      Period 、 BWidth 、 +Width the “Universal” knob to select

-Width、+Duty、-Duty            Time measure parameter.

Choose the measurement,

after pressing the button, the

Add                                     measurement value will be

shown         in    the     lower

left of the screen

Return to the first page of auto

Return

measurement menu.

Table 2-38 Auto measure function4-Delay measure menu:

Option             Setting                         Instruction

Source    CH1-CH2                      Select any two input signal

source for Delay measure.

Type      Phase、FRR、FRF、FFR、 Press the “Type” button or turn

FFF、LRR、LRF、LFR、LFF the “Universal” knob to select

Delay measure parameter.

Choose the measurement,

after pressing the button, the

Add                                    measurement value will be

shown          in    the     lower

left of the screen

Return                                 Return to the first page of auto

measurement menu.

SDS1000CML+/SDS1000DL+         User Manual  67

SIGLENT

Table 2-39 All Measurement function menu:

Option       Setting      Instruction

CH1        Select input signal channel.

Source

CH2

Turn on the all measurement function      to

On         measure voltage parameters.

Voltage

Off        Turn of the all measurement function      to

measure voltage parameters.

Turn on the all measurement function      to

On         measure Time parameters.

Time

Off        Turn of the all measurement function      to

measure Time parameters.

Return                   Return to the “All Measure main menu”.

Table 2-40 Introduction of the Measure Type:

Measure Type Introduction

The most positive peak voltage measured over the

Vmax             entire waveform.

The most negative peak voltage measured over the

Vmin             entire waveform.

Measures the absolute difference between the

Vpp              maximum and minimum peaks of the entire waveform.

Measures the highest voltage over the entire

Vtop             waveform.

Measures the lowest voltage over the entire waveform.

Vbase

Voltage between Vhig and Vlow of a waveform.

Vamp

The arithmetic mean over the first cycle in the

Vavg             waveform.

The arithmetic mean over the entire waveform.

Mean

The true Root Mean Square voltage over the first cycle

Crms             in the waveform.

The true Root Mean Square voltage over the entire

Vrms             waveform.

Defined as (Vmax-Vhig)/Vamp after the waveform

ROVShoot         rising.

Defined as (Vmin-Vlow)/Vamp after the waveform

FOVShoot         falling.

Defined as (Vmin-Vlow)/Vamp before the

RPREshoot        waveform rising.

Defined as (Vmax-Vhig)/Vamp before the

FPREshoot        waveform falling.

Rise Time Measures the time between 10% and 90%

Rise Time        of the first rising edge of the waveform.

68 SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

Fall Time Measures the time between 90% and 10% of

Fall Time the first falling edge of the waveform.

The duration of a burst.

BWid      Measured over the entire waveform.

+ Width Measures the time between the first rising

edge and the next falling edge at the waveform 50%

+ Wid

level.

-Width Measures the time between the first falling

edge and the next rising edge at the waveform 50%

- Wid

level.

Measures the first cycle waveform.

+ Duty    Positive Duty Cycle is the ratio between positive pulse

width and period.

Measures the first cycle waveform. Negative Duty

Cycle is the ratio between negative pulse width and

-Duty

period.

The amount one waveform leads or lags another in

time. Expressed in degrees, where 360 degrees

Phase

comprise one waveform cycle.

The time between the first rising edge of source X and

FRR       the first rising edge of source Y.

The time between the first rising edge of source X and

FRF      the first falling edge of source Y.

The time between the first falling edge of source X and

FFR       the first rising edge of source Y.

The time between the first falling edge of source X and

FFF       the first falling edge of source Y.

The time between the first rising edge of source 1 and

LRR      the last rising edge of source 2.

The time between the first rising edge of source X and

LRF      the last falling edge of source Y.

The time between the first falling edge of source X and

LFR      the last rising edge of source Y.

The time between the first falling edge of source X and

LFF       the last falling edge of source Y.

SDS1000CML+/SDS1000DL+     User Manual  69

SIGLENT

If you want to measure voltage parameters, please follow next steps:

1). Press the “MEASURE” button to enter the “Auto measurement” menu.

2). Press the top first option button to enter the “second measurement menu”.

3). Select measure type. If you press the “Voltage” option button, “Voltage

measurement” menu will display on the screen.

4). Press the “Source” option button to select “CH1”, “CH2” according to input

signal channel.

5). Press the “Type” option button to select the parameter type that you want

to measure, then press the “add” option button to add the measurement .

The corresponding icon and value will displayed    in the lower left of the

screen.

Picture 2.11-7

6). Press the “Return” option button to return the home page of “Auto

Measurement” menu. The selected parameter and the corresponding value

will display on the top first position of the home page.

You can display the other parameters and its’ value on the corresponding

position using the same way, and the screen can display five parameters at

one time.

70   SDS1000CML+/SDS1000DL+          User Manual

SIGLENT

If you want to measure time parameters using all measure function,

please follow next steps:

1). Press the “MEASURE” button to enter the “Auto Measure menu”.

2).Press the top option button to enter the second page of “Auto Measure

menu”.

3). Press “All Mea” option button to enter the “All Measure menu”.

4). Press “Source” option button to select input signal channel.

5). Press the “Time” option button to select “On”. Now all of time parameters

value will display on the screen at the same time.(See Picture)

Picture 2.11-8

SDS1000CML+/SDS1000DL+      User Manual 71

SIGLENT

2.12 Storage System

Showing as following picture, The SAVE/RECALL is the Storage System Function

Button.

Picture 2.12-1

You can quickly save and recall up to 20 oscilloscope panel settings, 10 groups’

waveforms in internal memory. There is a USB Host interface in the front panel of

the oscilloscope and you can save setup data, waveform data, waveform interface

image, CSV file to a USB flash drive furthest at a time. Setup data and waveform

data can be recalled to the current oscilloscope or the other same model

oscilloscopes. Picture data and CSV file can’t be recalled to the oscilloscope, but it

can be opened on the correlative computer software.

■ Using SAVE /RECALL Screen

The SAVE ALL screen is divided into functions based on Files or Directory.

Files shows option button for New File, Del File and Load.

Picture 2.12-2

72  SDS1000CML+/SDS1000DL+       User Manual

SIGLENT

While Directory shows option buttons for New Dir., Del Folder.

Picture 2.12-3

Recalling Files

The Load button is used to recall your setup files. Once you’ve navigated to the

desired file and it’s highlighted in the main screen area, press the Load option

button and the setup is recalled from the USB flash drive.

Note: The Load button option is disabled when BMP or CSV file types are

selected.

Both Directory and Files have Rename and Return option buttons on Page2/2.

Use these buttons to give an existing folder of file a new name or to exit the

SAVE/REC ALL screen, respectively.

Picture 2.12-4

SDS1000CML+/SDS1000DL+      User Manual  73

SIGLENT

Creating Directory and File

Create new directory and file by pressing the New Dir or New File option button.

The following screen is shown.

Picture 2.12-5

● The New Filemenu choices and behavior is the same as the New Folder menu.

It just has a different heading. The InputChar option button adds the selected

character to the courser position in Name field.

● Move the cursor position in the name field using the “→”and “←” option buttons.

● Turn the Universal knob to move through character selections. When the

desired character is highlighted, press the Universal knob or press he

“InputChar” option button to add it to the specific position in the Name field.

● Additional selections for BackSpace, DeleteCharacter, and CleanName are

available for convenience and are also accessed using the “Universal” knob in

the same manner.

● Press the “Confirm” option button (when you have the Name field completed as

desired) to save the file onto the memory device. After the “Confirm” option

button is pressed, a “Data Store Success!” messages briefly shown and the new

folder or file is shown on your USB flash drive.

74    SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

SAVE/RECALL SETUP

■ Save Setups to Device: The complete setup is stored in nonvolatile memory.

When you recall the setup, the oscilloscope will be in the mode from which the

setup was saved.

The oscilloscope saves the current setup if you wait three seconds after the last

change before you power off the oscilloscope. The oscilloscope recalls this setup

the next time you apply power.

Table 2-41 Save setup to device function menu:

Option       Setting     Introduction

Menu for the Storage/Recall setting in the

Type        Setups

oscilloscope

Save setup to the oscilloscope’s internal

Save to      Device

memorizer.

Press the “Setup” option button or turn the

Setup    No.1 to No.20

“universal” knob to select storage position.

Save                    Accomplish the storage

Recall                   Recall the storage in the “Setup” operation

Picture 2.12-6

Save setups to the oscilloscope’s internal memorizer, follow next steps:

For example: Save setup that set waveform display type to “Dots” to internal

memory.

1)Press the “SAVE/RECALL” button to enter the “SAVE/RECALL” menu.

SDS1000CML+/SDS1000DL+       User Manual    75

SIGLENT

2)Press the “Type” option button to select “setups”.

3)Press the “Save to” option button to select “Device”.

4)Press the “Setup” option button to select “No.1”.

5)Press the “DISPLAY” button to enter the “Display” menu.

6)Press the “Type” option button to select “Dots”.

7)Press the “SAVE/RECALL” button to enter the “SAVE/RECALL” menu.

8)Press the “Save” option button.

Recalling setup, follow next steps:

1)Press the “SAVE/RECALL” button to enter the “SAVE/RECALL” display menu.

2)Press the “Type” option button to select “Setups”.

3)Press the “Save to” option button to select “Device”.

4)Press the “Setup” option button or turn the “Universal” knob to select “No.1”.

5)Press the “Recall” option button.

76 SDS1000CML+/SDS1000DL+       User Manual

SIGLENT

■Save Setup to USB flash drive

Table 2-42 Save setup to USB flash drive function menu:

Option      Setting             Introduction

Type        Setups              Menu for the Storage/Recall settings.

Save to     File                Save setup data to USB flash drive.

Go to the Save/Recall interface and save

Save

the set file

Recall                          Recall the saved set file

Picture 2.12-7

Save setup to USB flash drive, follow these steps:

For example: Save setup that set waveform display type to “Dots” to USB flash

drive.

Press the “SAVE/RECALL” button to select “Setups”.

1) Insert USB flash drive to USB host port of the oscilloscope and wait that the

oscilloscope has initialized USB flash drive (about five second).

2) Press the “Save to” option button to select “file”.

3) Press the “Save” option button then you’ll go into the Save/Recall interface.

4) Press the “New folder” button to create a new folder

5) Press the “Del folder” option button to delete a folder

6) Press the “Modify” option button to modify a folder.

7) Press the “New file” option button to create a new file

8) Press the “Del file” option button to delete a file

9) Press “the next page” option button and press the “Rename” option button to

modify the name of the file or the folder.

SDS1000CML+/SDS1000DL+      User Manual   77

SIGLENT

10) Turn the “Universal” knob to select the folder and press the “Confirm” option

button and you save the set to the USB drive.

Recall setup data from USB flash drive, follow next steps:

1) Press the “SAVE/RECALL” button.

2) Press the “Type” button to select “Setups”.

3) Insert USB flash drive to USB host port of the oscilloscope and wait that the

oscilloscope has initialized USB flash drive (about five seconds).

4) Press the “Save to” option button to select “File”.

5) Press the “Save” option button then you’ll go into the Save/Recall interface .

6) Choose the file you want then press the “Load” option button (about five

seconds, there is a prompting message that “Read data success”), now setup

data have been recalled from USB flash drive.

78   SDS1000CML+/SDS1000DL+       User Manual

SIGLENT

■ RECALL FACTORY

You can use this option to recall the factory setup.

Table 2-43 Factory function menu:

Option             Setting           Instruction

Type              Factory           To view the Factory setup.

Change the bandwidth of        the

Update BW

oscilloscope.

Depth                               Recover the memory

Recover

Update cfg                             Update the config file

Load                                Recall the Factory setup.

Picture 2.12-8

SAVE/RECALL WAVEFORM

■Save waveforms to Device

Table 2-44 Save waveform to device function menu:

Option          Setup         Introduction

Menu for the Storage/Recall waveforms in

Type        waveforms

the oscilloscope.

Save waveforms to the oscilloscope’s

Save to         Device

internal memorizer.

Press the “waveform” option button or turn

waveform No.1 to No.10 the “Universal” knob to select storage

position.

Save                         Accomplish the storage.

Recall the storage in the “waveform”

Recall

operation

SDS1000CML+/SDS1000DL+      User Manual  79

SIGLENT

Picture 2.12-9

Save waveforms in internal memory, follow next steps:

1). Input a sine signal to channel 1 and press the “Auto” button.

2). Press the “SAVE/RECALL” button to enter “SAVE/RECALL” display menu.

3). Press the “Type” option button to select “waveforms”.

4). Press the “Save to” option button to select “Device”.

5). Press the “waveform” option button or turn the “Universal” knob to select

“No.1”.

6).Turn the “Volts/div” knob or “S/div” knob to adjust the waveform that you want

to need.

7). Press the “Save” option button.

Recall waveforms, follow next steps:

1). Press the “SAVE/RECALL” button to enter “SAVE/RECALL” display menu.

2). Press the “Type” option button to select “waveforms”.

3). Press the “Save to” option button to select “Device”.

4). Press the “waveform” option button or turn the “Universal” knob to select

“No.1”.

5). Press the “Recall” option button.

80 SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

■ Save waveforms to USB flash drive

Table 2-45 Save waveforms to USB flash drive function menu:

Option          Setup        Introduction

Type        Waveforms       Menu for the Storage/Recall waveforms.

Save to           File       Save waveforms to USB flash drive.

Save                        Accomplish the storage.

Picture 2.12-10

Save waveforms to USB flash drive, follow next steps:

1)Input a sine signal to channel 1, press the “AUTO” button.

2)Press the “SAVE/RECALL” button to enter the “SAVE/RECALL” display menu.

3)Press the “Type” option button to select “Waveforms”.

4)Insert USB flash drive to USB host port of the oscilloscope and wait that the

oscilloscope has initialized USB flash drive (about five seconds).

5)Press the “Save to” option button to select “File”.

6)Press the “Save” option button then you’ll go into the Save/Recall interface.

7)Create a file then press the “Confirm” button (about five seconds, there is a

prompting message that “Save data success” displaying on the screen), now

waveform data have been saved to USB flash drive

SDS1000CML+/SDS1000DL+       User Manual  81

SIGLENT

Recall waveforms form USB flash drive, follow next steps:

1). Press the “SAVE/RECALL” button.

2). Press the “Type” button to select “Waveforms”.

3). Insert USB flash drive to USB host port of the oscilloscope and wait that the

oscilloscope has initialized USB flash drive (about five seconds).

4). Press the “Save to” option button to select “File”.

5). Press the “Save” option button then go into the Save/Recall interface.

6). Choose the file you want then press the “Load” option button (about five

seconds, there is a prompting message that “Recall data success” displaying

on the screen), now waveform data have been recalled from USB flash drive.

■ Save Picture

Waveform interface image can be saved to USB flash drive, but they can’t be

recalled. You can open them on correlative computer software.

Table 2-46 Picture storage function menu:

Option      Setting       Introduction

Menu for the Storage/Recall        waveform

Type       Pictures

interface image.

Save                     Go to the Save/Recall interface.

Picture 2.12-11

82  SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

Save waveform image to USB flash drive, follow next steps:

1). Select the screen image that you want.

2). Press the “SAVE/RECALL” button to enter “SAVE/RECALL” menu.

3). Press the “Type” option button to select “Pictures”.

4). Insert USB flash drive to USB host port of the oscilloscope and wait that

the oscilloscope has initialized USB flash drive (about five seconds).

5). Press the “Print Key” option button to select “Save Picture”.

6). Press the “Save” option button then go into the Save/Recall interface”.

7). Press the “Save” button then you’ll go into the Save/Recall interface

8). Create the picture name then press the “Confirm” button (about five seconds,

there is a prompting message that “Save data success” displaying on the screen),

now picture data have been saved to USB flash drive.

Save/Recall CSV

Table 2-47 CSV storage function menu:

Option Setting         Introduction

Type      CSV          Menu for the Storage CSV file to USB flash drive.

Data      Displayed Set to store displayed waveform data to CSV file.

Depth     Maximum Set to store maximum waveform data to CSV file.

Para      On           Set whether store parameters to CSV file or not.

Save      Off

Save                   Go to the Save/Recall interface.

Picture 2.12-12

SDS1000CML+/SDS1000DL+       User Manual 83

SIGLENT

Save CSV file to USB flash drive, follow these steps:

1). Press the “SAVE/RECALL” button to enter “SAVE/RECALL” menu.

2). Press the “Type” option button to select “CSV”.

3). Insert USB flash drive to USB host port of the oscilloscope and wait that the

oscilloscope has initialized USB flash drive(about five seconds).

4). Press the “Data Depth” option button to select “Displayed” or “Maximum”.

5). Press the “Para Save” option button to select “On” or “Off”.

6). Press the “Save” option button then you’ll go into the Save/Recall

interface

7). Create a file name then press the “Confirm” button (About five seconds, there

is a prompting message that “Save data success” displaying on the screen),

now CSV file has been saved to USB flash drive.

84   SDS1000CML+/SDS1000DL+       User Manual

SIGLENT

2.13 Utility System

Picture 2.13-1

Table 2-48 Utility System function menu 1:

Option       Setting     Introduction

System                   Displays summaries of the oscilloscope

Status                   settings.

Open the key-press voice.

Sound

Close the key-press voice.

On          Turn on Frequency Counter

Counter

Off         Turn off Frequency Counter.

简体中文        Simplified Chinese

繁體中文        Traditional Chinese

English     English

‫العربیة‬     Arabic

Français    French

Deutsch     German

Language

Русский     Russian

Español     Spanish

Português Portuguese

日本語         Japanese

한국어         Korean

Italian     Italian

Next Page Page 1/4       Press this button to enter second page.

Picture 2.13-2

SDS1000CML+/SDS1000DL+     User Manual 85

SIGLENT

Table 2-49 Utility System function menu 2:

Option          Setting        Introduction

Do self cal                      Auto self emendation.

Screen Test      Run the screen detect program

Do Self

Keyboard Test       Run the keyboard detect program

Test

LED Test        Run the dot lighten detect program.

The oscilloscope connects to the

computer using USB cable. When you

Back USB         USBTMC          execute EasyScopeX software, please

select “USBTMC”. Now the computer

icon displays on the screen.

Next Page        Page 2/4        Press this button to enter third page.

Picture 2.13-3

Table 2-50 Functional Menu of the Utility System 3:

Option       Setting Introduction

Update                 You can update the oscilloscope use USB

firmware                flash Driver (About need two minutes).

Pass/Fail               Pres this button to enter the “Pass/Fail menu”.

Press this button to enter the waveform record

Record

menu.

IP Setting

IO Setting

GPIB

Next Page Page 3/4 Press this button to enter first page.

86 SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

Picture 2.13-4

Table 2-52 Function Menu of the Utility System :

Option         Setting                      Introduction

1min, 2min, 5min, 10min

Screen-saver 15min, 30min, 1hour, Set the time of Screen-saver

2hour, 5hour ,Off

Recorder                                    Enter recorder function

Next Page

Picture 2.13-5

SDS1000CML+/SDS1000DL+      User Manual 87

SIGLENT

2.13.1 System Status

Press the “System Status” option button of the Utility Menu to view the

oscilloscope’s hardware and software configuration.

Picture2.13-6

Table 2-53 System status instruction:

Option                    Introduction

Startup Times             List the boot-strap times.

Software version          List the software version.

Hardware Version          List the hardware version.

Product      type         List the product type.

Serial No.                List the product serial number.

2.13.2 Language

The oscilloscopes have twelve languages’ user menu to be selected.

Press the “Utility” button →“language” to select language.

88  SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

Picture 2.13-7

2.13.3 Self Calibration

Self Calibration procedure can optimize the signal path at the most measurement

precision. You can run this procedure at any time. If the operating temperature

changes by more than 5° C or the units run more than thirty minutes, you should

do the self calibration.

When you run the self Calibration, you should cut off all the probes and lead.

Then press the “Utility” button to choose the “Do self cal” to show the self

Calibration Menu, and execute self calibration program according to the

prompting on the screen.

Picture 2.13-8

SDS1000CML+/SDS1000DL+        User Manual  89

SIGLENT

2.13.4 Self Test

Press “UTILITY”→ “Do Self Test”

Table2-54

Option                    Introduction

Screen Test               Run Screen Test Program.

Keyboard Test             Run Keyboard Test Program.

LED Test                  Run LED Test Program.

Operation Steps:

1. Screen Test:

Select “Screen Test” to enter the screen test interface. The clew words “Press

‘SINGLE’ Key to continue,Press ‘RUN/STOP’ Key to exit” is displayed, you could

press the “Single” for test.

Picture 2.13-9

2. Keyboard Test

Select “keyboard Test” to enter the key test interface, the on-screen lathy

rectangle shapes represent the front panel keys. The shapes with two arrows

beside them represent the front panel knobs. The squares represent the knob

presses for scale knobs. Test all keys and knobs and you should also verify that

all the backlit buttons illuminate correctly.

Note:

●When you operate, the screen would display the white (Color LCD) and black

90   SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

(single LCD) repose.

●The tested button or knobs corresponding area would display green (Color LCD)

or white (single LCD).

●At the bottom of the screen display “Press‘ RUN/STOP’ Key Three Times to

exit” information prompt to show that press “RUN/STOP” three times for quitting

the test .

Picture 2.13-10

3. LED test

Select “LED Test” to enter the lighten interface, the on-screen lathy rectangle

shapes represent the front panel keys; the lathy rectangle shapes represent the

front panel knobs. The squares represent the knob presses for scale knobs. The

clew words “Press ‘SINGLE’ Key to continue , Press ‘RUN/STOP’ Key to exit” is

displayed, you could press the “Single” button continuously for testing, when

buttons are lighted ,the corresponding area on the screen would display

green(Color LCD) or white(Single LCD ).

Picture 2.13-11

SDS1000CML+/SDS1000DL+     User Manual   91

SIGLENT

2.13.5 Updating the System Software

■ Using USB flash drive update firmware

The software of the oscilloscope can be updated directly via USB Flash drive, this

process needs about two minutes. Follow next steps:

1. Insert USB Flash Drive with firmware procedure to USB Host interface on

the front panel of the oscilloscope.

2. Press the “UTILITY” button to enter the “Utility menu”.

3. Press “Next Page” option button to enter the third page of “Utility menu”.

4. Press the “Update Firmware” option button.

5. Press the “SINGLE” button to start updating the software according to prompt

on the screen.

Turn off the oscilloscope and turn on it again, the software has been updated. The

oscilloscope needs to execute “Do Self Cal” after update.

Note: Don’t cut off the power while the oscilloscope is updating.

2.13.6 Pass/Fail

The Pass/Fail function can monitor changes of signals and output pass or fail

signals by judging the input signal is within the predefined mask or not.

Table 2-55 Pass/Fail function menu 1:

Option          Setting       Instruction

Enable          On            Turn on Pass/Fail function.

Off           Turn off Pass/Fail function.

Source          CH1           Select signal input channel.

CH2

Operate                       Press to run the Pass/Fail test.

Press to stop the Pass/Fail test.

Msg Display On                Turn on times information display of the

Off           waveform pass/fail.

Turn off times information display of the

waveform pass/fail.

Next Page       Page 1/2      Press this button to enter second page of

Pass/Fail menu.

92   SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

Picture 2.13-12

Table 2-56 Pass/Fail function menu 2:

Option       Setting Instruction

On      Stop test when output occur.

Stop On Output               Continue test when output occur.

Off

Mask Setting                Enter the “Mask Setting menu”.

Return                   Return to the Pass/Fail main menu.

Next Page     Page 2/2 Return to the first page of Pass/Fail menu.

Picture 2.13-13

Table 2-57 Mask Setting function menu 1:

Option      Setting Instruction

X Mask                 Turn the “Universal” knob to set the range of

horizontal clearance to the waveform.

xdiv               <0.04div-4.00div>

Y Mask                 Turn the “Universal” knob to set the range of

ydiv                vertical clearance to the waveform.

SDS1000CML+/SDS1000DL+       User Manual 93

SIGLENT

<0.04div-4.00div>

Create                 Create a test mask according to the above

Mask                  clearance.

Internal

Location                Select position to store created mask.

External

Enter the second page of “Mask Setting

Next Page    Page 1/2

menu”.

Picture2.13-14

Table 2-58 Mask Setting function menu 2:

Option     Setting Instruction

Save                 Store created mask setting.

Load                 Recall stored mask setting.

Return                Return to the main menu of mask setting.

Last Page   Page 2/2    Return to the first page of “Mask Setting menu”.

Picture 2.13-15

94 SDS1000CML+/SDS1000DL+      User Manual

SIGLENT

RUN Pass/Fail test, please follow next page:

1)  Press UTILITY button to enter the “Utility menu”.

2)  Press the “Next Page Page1/4” option button.

3)  Press the “Next Page Page2/4” option button to enter the third page of “Utility

menu”

4)  Press the “Pass/Fail” option button to enter the “Pass/Fail” function menu.

5)  Press the “Enable Test” option button to select “On”.

6)  Press the “Source” option button to select input signal channel.

7)  Press the “Next Page Page1/2”option button to enter the second page of the

“Pass/Fail menu”.

8)  Press the “Mask Setting” option button to enter the first page of “Mask menu”.

9)  Press the “X Mask” button; turn the “Universal” knob to adjust the horizontal

clearance.

10)  Press the “Y Mask” button; turn the “Universal” knob to adjust the vertical

clearance.

11)  Press the “Create Mask” button to create mask and you also can enter the

next page of “Mask menu” to recall the stored mask.

12)  Enter the second page of “Pass/Fail function menu”, press the “Output”

option button to set the output option.

13) Enter the first page of “Pass/Fail function menu”, press the “Operate” option

button to select “ ”to run the pass/fail test.

Picture 2.13-16

SDS1000CML+/SDS1000DL+        User Manual  95

SIGLENT

2.13.7 Waveform Record

Waveform record can record input waveform from CH1 and CH2 with a maximum

record length of 2500 frames. This record behavior can also be activated by the

pass/fail test output, which makes this function especially useful to capture

abnormal signals in long term without keeping an eye watching it.

Waveform recorder: Record the waveforms with specified interval.

Table 2-59 Waveform recorder function menu:

Option        Setting            Instruction

Record             Set Record function menu.

Play Back          Set Play Back function menu.

Mode

Storage            Set Storage function menu.

Off                Turn off waveform recorder menu.

CH1

Source        CH2                Select record source channel.

P/F-OUT

Set time interval between record

Interval

frames.

End Frame                        Set max number of record frames.

(Record) Press to start recording.

Operate

(Stop)         Press to stop recording.

Picture2.13-17

96  SDS1000CML+/SDS1000DL+      User Manual

SIGLENT

To record waveforms, follow these steps:

1. Press the UTILITY button to enter “Utility menu”.

2. Press the “Next Page” option button to enter the third page of “Utility menu”.

3. Press the “Record” button to enter “Waveform Record Menu”.

4. Press the “Mode” option button to select “Record”.

5. Press the “Source” option button to select the input signal channel.

6. Select “Interval” option, turn the “Univeral” knob to adjust time interval between

record frames.

7. Select “End Frame”, select the “Universal” knob to adjust the max record

frame.

8. Press “ ”in “Operate” option to record waveform.

Play Back: Play back current record waveforms or saved record waveforms.

Table 2-60 Waveform play back function menu 1

Option         Setting       Instruction

Mode         Play Back      Set the Play Back function menu.

(Run)        Press to start playback playing.

Operate

(Stop) Press to stop playing.

Set circular play mode.

Play Mode

Set single time play mode.

Interval                     Set time interval between frames.

Next Page       Page 1/2      Enter the second page of Play Back menu.

Picture 2.13-18

SDS1000CML+/SDS1000DL+        User Manual   97

SIGLENT

Table 2-61 Waveform play back function menu 2:

Option         Setting       Instruction

Start Frame                    Set start frame.

Curr_Frame                     Select current frame to be played.

End Frame                      Set End frame.

Return                       Press to return the waveform

recorder main menu.

Next Page        Page 2/2      Return to the first page of Play

Back function menu.

Picture 2.13-19

To play back current record waveforms, follow these steps:

1. Press UTILITY button to enter “Utility menu”.

2. Press the “Mode” option button to select “Play Back”.

3. Press the “Play Mode” option button to select “      ” or “   ”.

4. Press the “Interval” option button to select time interval between play back

frames.

5. Press the “Next Page 1/2” to enter the second page of “Play Back function

menu”.

6. Select the “Start Frame” option, turn the “Universal” knob to adjust the start

frame of play back waveform.

7. Select the “End Frame” option, turn the “Universal” knob to adjust the end

frame of play back waveform.

8. Press the “Next Page 2/2” option button to return the first page of “Play Back

menu”.

9. Press the “ ” option button in “Operate” option to play back waveforms.

98  SDS1000CML+/SDS1000DL+      User Manual

SIGLENT

2.13.8 Recorder

The waveform recorder is a kind of seamless and no-gap real time recording of

waveform, means oscilloscope can save and replay waveform every time it

captured. It is similar to waveform recording instrument, and the biggest recording

size of its internal memory is 6M

Table 2-62 waveform recording function menu:

option                               description

Record          record the waveform persistently

Replay          Replay recorded waveform, refer to table 2-64

Option         Set the recorder properly, refer to table 2-65

Return         Quit the recording function

Picture 2.13-20

Table 2-63 waveform replay function menu

Option                               Description

Pause          Pause the replay function

Restart         Restart the replay function

Previous         Recall the waveform again

Next          Fast forward play the recalled waveform again

Return         Quit the recalling interface

SDS1000CML+/SDS1000DL+         User Manual 99

SIGLENT

Picture 2.13-21

Table 2-64 recorder’s setting menu

Option          Setting          Description

Full Screen      Waveform of Full screen recording

and channel recalling

Waveform of recording on split screen

Viewer

Split            and channel recalling, CH1 is

displayed in upper half-screen, CH2 is

displayed in lower half-screen.

Recorder      saves      the    channel

Continuous

waveform persistently; the latter

Record Mode                      waveform will cover the former.

Single

Recorder will stop waveform saving

when its memory is up to 7M.

By point         when recalling, the screen waveform

refresh from left to right

Replay Mode     By frame         When recalling, screen waveform will

refresh completely according to the

capturing time of each frame

Return                           quit the recording setting interface

Picture 2.13-22

100  SDS1000CML+/SDS1000DL+      User Manual

SIGLENT

Recorder Operation step:

1. Press the UTILITY button to enter “Utility menu”.

2. Press the “Next Page” option button to enter the four page of “Utility menu.

3. Press the “Recorder” button to enter the Recorder manual.

4. Press the “Option” button to select the parameter which you want by yourself

5. Press the “Recorder” button to enter the Record menu and then Press the

“Start” button to start record waveform

6. After finished record waveform, then Press the “Replay” button to view the

recorded waveform.

SDS1000CML+/SDS1000DL+    User Manual    101

SIGLENT

2.13.9 Remote Control

There are two means to remotely control the oscilloscope:

Use user-defined programming

User can control the oscilloscope by programming through standard SCPI

(Standard Commands for Programmable Instruments) commands. For more

details about commands and programming, Please refer to Programming Manual.

Use PC software offered by SIGLENT or other manufacturers

User can send command to remotely control the oscilloscope through PC software

EsayScope which offered by SIGLENT. Besides, “Measurement & Automation

Explorer” which provided by NI (National Instruments Corporation) also can be

used to control the oscilloscope. This oscilloscope can communicate with a PC by

a USB bus. This part will show how to use EsayScope to remotely control an

oscilloscope by USB interface.

Control by USB

Connect Instrument

Use a USB data wire to connect the oscilloscope (USB Device) to a PC.

Installing USB Device

Set oscilloscope Back USB to USBTMC. If you have installed EsayScope or NI

software, the PC will pop-up a hardware update guide dialogue box which shows

below when you connect the oscilloscope to the computer correctly for the first

time, and please install the “USB Test and Measurement Device” device program

follow the prompting messages. Below are the steps:

102   SDS1000CML+/SDS1000DL+      User Manual

SIGLENT

SDS1000CML+/SDS1000DL+ User Manual 103

SIGLENT

104 SDS1000CML+/SDS1000DL+ User Manual

SIGLENT

SDS1000CML+/SDS1000DL+ User Manual 105

SIGLENT

Research Device

Open EsayScope software; click “add device” to research, there will pop-up the

following dialogue box, then click Add.

Check Instrument Resource

The instrument information which has been researched is as the picture below. It

shows the instrument number and the USB interface information.

Communicate test

Click “SCPI control” and choose “common SCPI”: *IDN?, click Enter to send

command. The software will accept the data and display the instrument

information. See picture below.

106  SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

2.14 Online Help Function

The oscilloscope has an online help function that supplies multi-language help

information, and you can recall them to help you operate the oscilloscope when

you need.

Press the “HELP” button to enter the help status, then press every button to recall

the corresponding help information. Note:If you want to see help information of

SINGLE button and RUN/STOP button, you need to see them when you first enter

the help status because you press the SINGLE button can enter next page of help

information and Press the RUN/STOP button can return to the first page of help

information when other button’s help information exceed one page.

All submenus of every main menu have their help information. Note: If you want to

see help information of next page submenus, first you need to press the “HELP”

button to exit the help status and switch to next page menu then press the “HELP”

knob to enter the help status again and press the submenu option buttons to see

its help information.

Picture 2.14-1

SDS1000CML+/SDS1000DL+      User Manual  107

SIGLENT

Chapter 3 Prompting Messages and

Troubleshooting

3.1 Prompting Messages:

■ Trig level at limit! : Mention you that the trigger Level is at a limit when you turn

the Trig level knob.

■ Horizon position at limit! : Mention you that the horizontal position is at a limit

when you turn the horizon position knob.

■ Volts/Div at limit! : Mention you that the vertical voltage have already touched

the Min 2MV/div or the Max 10V/div.

■ Volts position at limit! : The system would display this information when the

vertical position is at a limit.

■ Sec/Div at limit! : Prompts that the Volts/Div is at full range while turning the

vertical scale knob.

■ Holdoff time at limit! : Turn the “Universal” knob when holdoff time has been to

max or min value, now the system will clew this information.

■ Functions isn’t useable! : Under several special modes, the some functions

could not be running.

■ No signal! : The system would clew this information when the signal could not

match the auto set condition. (Using in the auto set)

■ Adjust at limit! : You could adjust the pulse width by the universal knob till the

pulse width has reached min20.0ns or max 10.0s.

■ Location Empty! : If you have no stored waveforms or setups on some location,

the screen will display this information when you press the “Recall” button on

this location.

■ USB Flash Drive Plug In! : This information will appear when you invert the

USB Flash Drive to the USB Host port.

■ USB Flash Drive Pull Out! : This information will appear when you pull out the

USB Flash Drive.

■ Store Data Success! : Save setup data, waveform data or picture data to the

internal of the oscilloscope or USB flash successful.

108  SDS1000CML+/SDS1000DL+        User Manual

SIGLENT

■ Ready Data Success! : Read setup data or waveform data from the internal of

the oscilloscope or USB flash successful.

■ Please set Back USB to printer! : Press the “S/div” knob will appear this

information on the screen when the “Print Key” option is set to “Print Picture”

and the “Back USB” option is set to “USBTMC”.

■ USB Flash Drive isn’t connected! : When the “Save To” option is set to “File”

or the “Print Key” option is set to “Save Picture” in “Save/Recall” menu , Press

the “Save” option button or the “S/div” knob before inverting the USB Flash

Drive to the USB Host port will appear this information on the screen.

■ Record Wave Success! :This message will appear when you finish recording

waveforms.

SDS1000CML+/SDS1000DL+     User Manual  109

SIGLENT

3.2 Troubleshooting

1. After the oscilloscope is powered on,the screen remains dark ,please do

as following steps:

(1) Check the power cable’s connection.

(2) Ensure the power switch is turned on.

(3) After the inspections above,restart the oscilloscope.

(4) If the oscilloscope is still not used after the checking, please connect with

Siglent

2. If there is no signal wave in the screen after gathering the signal, please

do as following steps:

(1) Check the probe connecting with the signal cable or not

(2) Check the signal cable connecting with the BNC connector or not.

(3) Check the probe whether connect with the goods tested or not.

(4) Check the tested goods produce the signal or not.

(5) Gather the signal again.

3. The value of the tested voltage is 10 times higher/lower than the real one ,

please do as following steps:

Check the attenuation quotient whether match the probe attenuation proportion or

not.

4. Display the wave, but not steady, please do the following steps:

(1) Check the signal source on the trigger interface whether or not matches the

signal channel.

(2) Check the trigger mode: normal signal should use the “edge” trigger mode.

The video signal should use the “Video” Trigger mode. The signal would

display steady, only using the matching trigger mode.

(3) Attempt to change the “coupling” into “HF Reject” or “LF Reject” display, so that

the High/low frequency noise disturb the trigger should be filtrated

5. Press “RUN/STOP” button, but no display.

Check the trigger mode on the trigger interface whether or not in the “normal” or

“single”, and check the trigger level is whether or not over the wave range. If yes,

please put the trigger level to the middle position or set the trigger mode to the

“Auto” position. In another hand, you could choose the “Auto” button to set up

110 SDS1000CML+/SDS1000DL+         User Manual

SIGLENT

automatically.

6. After the Acquisition is set to Averages or Display Persistence time is set

too long, the waveform refreshes slowly.

It is normal in these settings

7. The signal is displayed as ladder like waveform

(1)This phenomenon is normal. The time base maybe is too slow .you should turn

the horizontal scale knob to increase horizontal resolution to improve the

display.

(2) Maybe the display Type is set to “Vectors”, You could set it to Dots mode to

improve the display.

SDS1000CML+/SDS1000DL+       User Manual  111

SIGLENT

Chapter 4 Service and Support

4.1 Maintain Summary

Siglent warrants that the products that it manufactures and sells will be free from

defects in materials and workmanship for a period of three (3) years (accessories

for a period of one year) from the date of shipment from an authorized Siglent’s

distributor. If a product or CRT proves defective within the respective period,

Siglent will provide repair or replacement as described in the complete warranty

statement.

To arrange for service or obtain a copy of the complete warranty statement,

please contact your nearest Siglent’s sales and service office.

Except as provided in this summary or the applicable warranty Statement, Siglent

makes no warranty of any kind, express or implied, including without limitation the

implied warranties of merchantability and fitness for a particular purpose. In no

Event shall Siglent be liable for indirect, special or Consequential damages

112 SDS1000CML+/SDS1000DL+          User Manual

SIGLENT

4.2 Contact SIGLENT

MTR Add:3//F, Bldg No.4, Antongda Industrial Zone, 3rd Liuxian Road, Bao’an

District, Shenzhen, 518101, P.R.China

Tel:0086-755-3661 5186

E-mail:sales@siglent.com

http://www.siglent.com

SDS1000CML+/SDS1000DL+   User Manual 113

Appendix A: Default Setup

options 、 knobs or

Menu or system                         Default setup

buttons

Coupling             DC

BW limit             Off

Volts/div            Coarse

CH1,CH2           Probe                1X

Invert               Off

Filter               Off

Volts/div            1.00V

Operation            CH1+CH2

CH1 Invert           Off

CH2 Invert           Off

FFT operation:

MATH              Source               CH1

Window               Hanning

FFT Zoom             1X

Scale                dBVrms

Display              Split

Window               Main

Position             0.00μs

HORIZONTAL        Sec/div              500μs

Window Zone          50.0μs

Trigger knob         level

Type                 Off

Source               CH1

CURSOR

Horizontal (voltage) +/-3.2divs

Vertical (time)      +/-5divs

three mode options   Sampling

Averages             16

ACQUIRE

Sampling mode        Real Time

Type                 Vectors

Persist              off

DISPLAY           Gird

Intensity            60%

Brightness           40%

Format               YT

114 SDS1000CML+/SDS1000DL+   User Manual

Menu Display          infinite

Type                  Setups

SAVE/RECALL     Save To               Device

Setup                 No.1

REFA/REFB             REFA

Source                CH1

REF

REFA                  off

REFB                  off

Sound                 on

Counter               On

UTILITY         Back USB              USBTMC

Pass/Fail             off

Record                off

Type                  edge

Source                CH1

Slope                 Rising

TRIGGER (edge)

Mode                  Auto

Coupling              DC

LEVEL                 0.00V

Type                  pulse

Source                CH1

When                  =

TRIGGER (pulse)

Set Pulse Width       1.00ms

Mode                  Auto

Coupling              DC

Type                  Video

Source                CH1

Polarity              Normal

TRIGGER (Video)

Sync                  All Lines

Standard              NTSC

Mode                  Auto

Type                  Slope

TRIGGER         Source                CH1

(Slope)         Time                  1.00ms

Mode                  Auto

Type                  Alternative

TRIGGER         Source                CH1

(Alternative)   Mode                  Edge

Coupling              DC

SDS1000CML+/SDS1000DL+  User Manual 115

Appendix B: Daily Maintain and Cleaning

Daily Maintain

DO not store or leave the instrument in where the LCD display will be

exposed to direct sunlight for long periods of time.

CAUTION: To avoid damage to the instrument or probes, do not expose

them to sprays, liquids, or solvents

Cleaning

If this instrument requires cleaning, disconnect it from all power sources

and clean it with a mild detergent and water. Make sure the instrument is

completely dry before reconnecting it to a power source.

To clean the exterior surface, perform the following steps:

1). Remove loose dust on the outside of the instrument and probes with a

lint-free cloth. Use care to avoid scratching the clear plastic display

filter.

2). Use a soft cloth dampened with water to clean the instrument. Use an

aqueous solution of 75% isopropyl alcohol for more efficient cleaning.

Note:

To avoid damage to the surface of the instrument or probes, do not use

any abrasive or chemical cleaning agents.

116 SDS1000CML+/SDS1000DL+         User Manual